NeurOne

Lens and Goggle Assembly

Tooling Specification

NP-TOOL-LENS-001  Rev B  |  2026-05-10  |  Pre-Tooling Draft

References: CLAUDE.md §3 (Visual Stimulation) · §7.1 (Critical Design Changes) · §7.2 (Locked Changes) · §8.1 (Service Network) · §13.4 (Pending Decisions)

1.  PURPOSE AND SCOPE

This document consolidates all molded geometry requirements, optical coating process specifications, and electrical contact specifications for the NeurOne lens rim assembly, goggle arm assembly, and snap-on shade assembly. Five locked design changes (sliding rail lens mount, N42 magnet positions, AgNW outer conductive coating, hard coat, and EC driver contact positions) each require co-designed features in two or more mating tooling packages. This document exists because the coupling between lens rim tooling, goggle arm tooling, and shade assembly tooling creates an omission risk: any feature missed before first-cut requires tooling modification across multiple packages (typically $15,000–40,000 per package and 6–8 week delay). All features below are MANDATORY and must appear in all relevant mold design reviews before steel is cut.

Scope: Lens rim body (thermoplastic or CFRP-composite), goggle arm body, snap-on shade assembly body, and optical coating process specifications for the standard lens and EC (electrochromic) premium lens. Not in scope: LED FPC assembly within the lens (covered by NP-HW-FPC-001), hub electrical interfaces, EEG cable routing (NP-DRV-SHELL-001).

Companion documents: NP-HW-FPC-001 Rev D (zone module FPC — LED array in lens), NP-TOOL-ZM-001 Rev A (zone module body tooling), NP-FAI-ZM-001 Rev A (zone module FAI), NP-COORD-001 Rev A (engineering coordination gates).

2.  FEATURE OVERVIEW — MANDATORY CHECKLIST

The following table lists all features required across lens rim, goggle arm, and shade assembly tooling. Every feature must appear in the respective mold design review before steel is cut. The mold design review gate is NP-COORD-001 G2 (Pre-Tooling). Any OPEN item blocks tooling release for the package in which it appears.

| ID | Feature | Tooling Package(s) | BOM Delta | Risk / Reference | Status |
| --- | --- | --- | --- | --- | --- |
| F-01 | Sliding rail groove — T-slot profile in lens rim inner face | Lens rim | +$1.20/lens | CLAUDE.md §7.1 | MANDATORY |
| F-02 | Sliding rail tongue — mating protrusion in goggle arm | Goggle arm | Included F-01 | CLAUDE.md §7.1 | MANDATORY |
| F-03 | N42 magnet pockets — 6× per lens rim, ≥1 mm polymer wall all faces | Lens rim | −$0.80 vs N52 | CLAUDE.md §7.2 | MANDATORY |
| F-04 | Shade mating magnet pockets — 6× per shade body, aligned to F-03 | Shade assembly (S1/S2/S3) | Included F-03 | CLAUDE.md §3 | MANDATORY |
| F-05 | EC driver contact pad positions — 2× per lens rim (EC lens only) | Lens rim | Included F-06 | CLAUDE.md §3 | MANDATORY — EC lens |
| F-06 | EC driver spring contact positions — 2× per goggle arm (EC lens only) | Goggle arm | +$1.80/lens est. | CLAUDE.md §3 | MANDATORY — EC lens |
| P-01 | AgNW outer conductive coating — replaces ITO on outer lens face | Lens optical stack (process) | +$8–12/lens | CLAUDE.md §7.1 | MANDATORY |
| P-02 | Hard coat — 3–5 µm silicone over AgNW and EC film surfaces | Lens optical stack (process) | +$4.00/lens | CLAUDE.md §7.2 | MANDATORY |
| P-03 | Inner PDMS diffuser — plasma-activated anti-fouling, scalp/eye-facing | Lens optical stack (process) | Minimal | CLAUDE.md §3 | MANDATORY |
| F-07 | Lens rim guard anchor hooks — tether attachment points on goggle arm | Goggle arm | +$0 (tooling) | CLAUDE.md §8.3 | MANDATORY |

3.  FEATURE SPECIFICATIONS

3.1  F-01 / F-02 — Sliding Rail Lens Mount

The sliding rail system replaces the legacy press-fit lens installation that previously required an alignment jig at Tier A optical service centres (NP-COORD-001, CLAUDE.md §8.1). The rail is a two-part mating geometry: a T-slot groove molded into the inner face of the lens rim (F-01) engages a corresponding tongue molded into the goggle arm (F-02). The lens is inserted from above and slides down until the spring detent clicks. No tooling, adhesive, or alignment jig is required. BOM delta: +$1.20/lens total for both mated parts.

Rail geometry requirements:

| Parameter | Specification | Rationale |
| --- | --- | --- |
| Rail profile | T-slot, 2.5 mm slot width at mouth, 4.0 mm slot width at base, 3.0 mm depth | Dovetail-style profile for positive retention; cannot pull out radially once seated |
| Rail length | Full width of lens rim at goggle arm interface; nominal 38 mm (confirm from lens CAD) | Distributes insertion load across full interface width |
| Insertion direction | Superior-to-inferior (top of rim first, slide downward into goggle arm) | Gravity-assisted seating; shade removal force does not oppose gravity |
| Rail tolerance | ±0.08 mm on slot mouth width; ±0.10 mm on base width | Reliable shade engagement across all three shade types (S1/S2/S3) |
| Tongue fit clearance | 0.05–0.10 mm bilateral clearance between tongue and groove flanks | Slip fit for user self-install; no interference force required |
| Retention detent | Spring-loaded PEEK ball detent at inferior end of rail travel; audible click at full engagement; extraction force ≥ 3 N in pull-out direction | Confirms engagement without visual inspection; ≥ 3 N prevents accidental disengagement during head movement |
| Extraction force | ≤ 5 N axial pull (superior direction) to disengage detent and remove lens | User self-service without tools; accessible to users with reduced hand strength |
| Shade compatibility | Common rail geometry on all three shade bodies (S1 opaque, S2 polarising, S3 prescription carrier); Rx insert engages shade carrier, not rail directly | Single tooling family; S3 Rx insert is a separate snap-fit into shade carrier |
| Material — lens rim | Lens rim body: glass-reinforced PA12 (30% GF) or CFRP composite; rail groove may be machined or molded; surface finish Ra ≤ 0.8 µm in groove flanks | PA12 GF provides dimensional stability under thermal cycling; smooth groove surface prevents galling |
| Material — goggle arm | Goggle arm body: same material family as lens rim or CFRP; tongue molded integral to arm body | Matched CTE prevents thermal differential causing bind or looseness across operating temperature range −10 to +45 °C |
| Cycling life | ≥ 500 lens insertion/extraction cycles without detent force degradation > 20% | User expected to swap S1/S2/S3 shades multiple times per week; 500 cycles ≈ 2 years at 5 swaps/week |

DRAWING REQUIREMENT: Rail groove cross-section with T-slot dimensions, tongue profile, and detent position must appear on lens rim and goggle arm detail drawings before mold review. Both drawings must be reviewed together — mating parts cannot be approved independently.

3.2  F-03 / F-04 — N42 Magnet Pockets

Six N42 neodymium magnets per lens rim provide the snap-on attachment force for all three shade types (S1, S2, S3). N42 grade is specified over N52: N42 offers better impact tolerance under corner drop events; N52 is brittle under concentrated impact. BOM delta: −$0.80 vs N52 (N42 lower material cost). The shade assembly carries six corresponding mating magnets (F-04) with opposing polarity pattern.

Magnet pocket requirements (F-03 — lens rim):

| Parameter | Specification | Rationale |
| --- | --- | --- |
| Magnet grade | N42 neodymium (NdFeB sintered); confirm final supplier grade from magnet datasheet | N42: Br ≈ 1.29 T, Hcj ≈ 1,195 kA/m. Better impact fracture resistance than N52 (Br ≈ 1.45 T) at modest magnetic performance cost. Shade snap force adequate at N42. |
| Magnet geometry | 5 mm diameter × 2 mm thick disc; axially magnetised | Standard stocked disc geometry; minimises tooling-side pocket complexity; axial magnetisation maximises normal snap force at ≤ 1 mm air gap |
| Count | 6 magnets per lens rim (12 total per headset) | Provides ≥ 800 g total normal pull per lens with N42 at 0.5 mm air gap; sufficient to retain shade during session movements |
| Pocket depth | 2.1 mm (magnet recessed 0.1 mm below lens rim outer face) | Slight recess protects magnet face from direct contact impact; magnet does not protrude |
| Polymer wall — all faces | ≥ 1.0 mm wall thickness on EVERY face (superior, inferior, medial, lateral, and outer face) | LOCKED REQUIREMENT (CLAUDE.md §7.2). N42 magnets cleave along field lines under point impact. ≥ 1 mm wall on outer face absorbs impact before reaching magnet. Verify in CAD cross-section — BLOCKING for mold approval. |
| Pocket tolerance | 5.05 + 0.00 / −0.03 mm diameter × 2.10 + 0.00 / −0.05 mm depth | Press-fit depth; adhesive bond (UV-cure optical adhesive, ≥ 5 MPa shear strength) provides primary retention; pocket geometry provides anti-rotation |
| Magnet adhesive | UV-cure cyanoacrylate or epoxy, ≥ 5 MPa lap shear on PA12 GF; no proud adhesive fillet | Adhesive must not intrude into shade mating gap; flush bond line only |
| Position layout | 3-2-1 pattern per lens: 2 lateral positions (medial and lateral pole of lens), 1 superior center; confirm geometry in CAD for equal shade-retention force distribution | Symmetric pull prevents shade rocking under lateral head acceleration; 3-2-1 layout standard for rectangular magnet frames |
| Polarity pattern | Alternating N-S-N (superior to inferior) to prevent shade mis-orientation; shade body magnets mirror this pattern | Prevents 180° rotational error on S1/S2 — shades are symmetric in appearance but optics asymmetric; single valid orientation enforced by magnet polarity |
| Shade compatibility | F-04 mating pockets in S1, S2, and S3 carrier bodies must share identical 6-position magnet pattern | Single pattern enables all shade types to mate with the same lens rim without re-tooling shades |
| Cycling durability | Snap/unsnap cycle life ≥ 10,000 cycles without >10% pull-force degradation | Magnet N42 is corrosion-resistant with Ni-Cu-Ni plating. Cycling degrades plating, not magnet performance. Confirm plating spec from supplier. |

DRAWING REQUIREMENT: Magnet pocket CAD cross-section must show ≥ 1 mm wall on ALL faces. ME sign-off on this cross-section is BLOCKING for mold design review. Pocket positions must be confirmed in CAD assembly against all three shade body CAD models before mold release.

MATING REQUIREMENT (F-04): Shade body tooling spec (separate document) must specify the identical 6-position magnet pocket pattern with reversed polarity sequence. F-03 and F-04 mold design reviews must be conducted concurrently by the same ME reviewer.

3.3  F-05 / F-06 — EC Driver Contact Positions (EC Lens Only)

The electrochromic (EC) premium lens (+$89 upgrade / $129 standalone) requires an electrical interface between the goggle arm and the lens rim to drive the EC film. The EC film is bistable: it holds its tint state indefinitely without power (~15 mW hold during transition only). Power is required only during state transitions (≤ 2 s). On power restore, the EC driver commands the lens to 75% VLT (safety failsafe — clear state). The EC driver monitors transition time as a proxy for contact resistance: transition time > 3 s flags a "clean contacts" prompt in the app. Contact positions must be specified in both lens rim tooling (F-05) and goggle arm tooling (F-06) before either is cut.

EC contact requirements (F-05 — lens rim pad positions):

| Parameter | Specification | Rationale |
| --- | --- | --- |
| Contact count | 2 per lens rim (positive and negative EC film electrodes) | EC film is a capacitive sandwich requiring two electrodes. Current path: goggle arm driver → spring pin (F-06) → rim pad (F-05) → EC film ITO/AgNW bus bar → EC active layer |
| Contact pad geometry | Ø 4.0 mm circular contact area; recessed 0.15 mm below lens rim inner face within a Ø 5.0 mm recess pocket | Recess prevents pad contamination from dust or sebum during lens swap; spring pin tip engages recess automatically on lens insertion |
| Plating specification | Hard gold ≥ 0.5 µm cobalt-alloyed over nickel underplate ≥ 2 µm on copper substrate | Consistent with hub power contact spec (CLAUDE.md §7.2). Hard gold resists fretting corrosion over repeated lens insertion cycles. Cobalt alloy improves hardness vs. pure gold. |
| Contact resistance (new) | ≤ 50 mΩ per contact pair at ≥ 150 g spring contact force | EC film switching current ≤ 50 mA during transition. At 50 mΩ, resistive drop ≤ 2.5 mV — negligible vs. EC drive voltage (typically 1.5–3 V). |
| Contact resistance (SHDR flag threshold) | > 200 mΩ triggers "clean contacts" app notification (via transition time >3 s) | EC driver measures transition time each switch event. Transition time proportional to series resistance at fixed drive voltage. >3 s transition corresponds to >200 mΩ total path resistance at nominal EC drive voltage. |
| Position — contact 1 (positive) | Superior-medial corner of lens rim inner face; exact coordinates from CAD; ≥ 8 mm from nearest magnet pocket (F-03) | ≥ 8 mm separation prevents fringe field from N42 magnet from affecting contact spring positioning in goggle arm |
| Position — contact 2 (negative) | Superior-lateral corner of lens rim inner face; symmetric to contact 1 about lens optical axis; ≥ 8 mm from nearest magnet pocket | Symmetric placement equalises contact force when lens is seated; lateral-to-medial asymmetry provides polarity keying with matching goggle arm pattern |
| Standard lens compatibility | F-05 pockets present in all lens rim tooling (standard and EC). Standard lens rim: pockets are cosmetic blind features — no conductive pad installed. EC lens rim: pads installed. | Single lens rim mold serves both standard and EC lens; eliminates second mold. Pad is populated at assembly for EC configuration only. |
| Lens orientation key | Contact positions are asymmetric L-R: prevents EC lens from being installed rotated 180°; lens rim guard and shade rail also asymmetric for same reason | EC film has a defined current polarity — reversing polarity degrades EC film. Three independent keying mechanisms (contacts, rail detent position, lens rim guard hook) each independently prevent reversal. |

EC contact requirements (F-06 — goggle arm spring pins):

| Parameter | Specification | Rationale |
| --- | --- | --- |
| Spring pin type | Pogo-pin style: BeCu spring barrel with hard-gold-plated tip; Ø 1.4 mm barrel, Ø 1.0 mm tip; stroke 1.5 mm; spring force 150–350 g at mid-stroke | BeCu spring maintains contact force over >10,000 compression cycles without set. Hard gold tip matches rim pad plating for galvanic compatibility. 150–350 g provides reliable contact without excessive lens insertion resistance. |
| Count | 2 per goggle arm (one per EC electrode) | Matches F-05 pad count on lens rim |
| Mounting | Press-fit into blind bores in goggle arm body; bore diameter 1.42 +0.00/−0.02 mm; bore depth 8 mm | Press-fit retention provides >50 N axial pull-out resistance. Blind bore prevents liquid ingress to goggle arm interior. |
| Spring pin protrusion | 0.8 mm proud of goggle arm inner face at mid-stroke when lens is seated | Ensures >150 g contact force is maintained across ±0.15 mm lens-to-arm positional tolerance from sliding rail |
| Electrical connection | Pogo barrel connects to EC driver PCB via 28 AWG wire or flex; wire routed within goggle arm internal channel | EC drive current ≤ 50 mA; 28 AWG rated 0.36 A continuous — adequate with >7× margin |
| Position tolerance | ±0.10 mm on pogo pin bore position relative to lens rim contact pad position (in mated assembly) | Ensures tip engages Ø 4.0 mm pad with ≥ 1.0 mm margin even at worst-case tolerance stack-up from sliding rail clearance |
| SHDR logging | EC driver logs transition time per switch event to SHDR as contact resistance proxy. Trend over 30 days flags cleaning need before hard failure. | Contact resistance trend is device health data only (no user biology) → SHDR. Enables fleet predictive maintenance for EC contact wear. |

DRAWING REQUIREMENT: EC contact pad positions (F-05) and spring pin bore positions (F-06) must appear on lens rim and goggle arm drawings with positional tolerances. Both drawings must be reviewed together. Contact polarity (positive/negative) must be marked on lens rim drawing and goggle arm drawing.

EC lens only: Features F-05 and F-06 are required only when the EC lens configuration is ordered. Standard lens uses the same lens rim mold but F-05 pads are not populated. Goggle arm tooling is identical for both; spring pins are populated for EC configuration only.

3.4  P-01 — AgNW Outer Conductive Coating (Replaces ITO)

The outer face of both standard and EC lenses carries a conductive coating. ITO (indium tin oxide) is the conventional choice but is disqualified: ITO has 0.5% strain-to-failure and cracks under flex or point impact (relevant to lens rim snap-in and drop events). Silver nanowire (AgNW) network replaces ITO: AgNW tolerates 5–10% strain without conduction loss, is compatible with the overlying hard coat (P-02), and maintains 85–90% optical transmission. BOM delta: +$8–12/lens over standard uncoated lens.

AgNW coating specification:

| Parameter | Specification | Rationale |
| --- | --- | --- |
| Coating material | Silver nanowire (AgNW) percolating network; wire diameter 20–40 nm, length 10–30 µm; embedded in polymer matrix (e.g., PEDOT:PSS or proprietary binder) | AgNW diameter and aspect ratio determine network percolation density and optical haze. Specify from supplier datasheet — final wire geometry must be confirmed from supplier COA. |
| Sheet resistance | 20–30 Ω/sq on 380 µm PC substrate at 85% transmission | Adequate for EC film drive bus bar function on EC lens. On standard lens, AgNW provides EMI shielding (incidental benefit) and anti-static surface to prevent dust attraction to lens surface. |
| Optical transmission | ≥ 85% at 550 nm (photopic peak); ≥ 82% at 660 nm and 808 nm (PBM therapeutic wavelengths) | Transmission at 660 nm and 808 nm is critical — NeurOne Mode F delivers 808–830 nm retinal PBM through the lens. AgNW must not attenuate PBM dose below therapeutic threshold. Verify on representative coupon at both wavelengths before production release. |
| Optical haze | ≤ 1.5% (ASTM D1003) | Haze above 1.5% causes visible scatter during visual stimulation protocols (EMDR, photic driving). AgNW network geometry optimized for low haze by supplier. |
| Substrate | Applied to outer face of polycarbonate (PC) lens; PC substrate ≥ 380 µm thickness; primer layer per supplier recommendation before AgNW deposition | Primer ensures adhesion of AgNW network to PC. Specify primer from AgNW supplier datasheet — generic primers may cause delamination under thermal cycling. |
| Flex performance | No resistance increase >20% after 1,000 bend cycles at R = 50 mm (IEC 62368-1 flex test modified) | Lens rim snap-in introduces localized flex at lens edge. 50 mm radius is conservative for lens geometry. 1,000 cycles = expected swap life. |
| Strain-to-failure | ≥ 5% elongation before electrical open (vs. ITO ≤ 0.5%) | LOCKED DECISION (CLAUDE.md §7.1). 5% minimum confirms AgNW meets specification. Point drop impact on lens center estimated at <1% local strain for 380 µm PC at 1 m drop — 5× margin. |
| Environmental stability | Sheet resistance stable within ±20% after 500 h at 85°C / 85% RH (IEC 60068-2-78) | AgNW networks can silver-migrate under humidity. Require supplier qualification data for 85/85 test on proposed AgNW / binder / hard coat stack. Hard coat (P-02) provides primary moisture barrier — test must be performed on full P-01 + P-02 stack. |
| Hard coat compatibility | AgNW coating must accept P-02 hard coat adhesion; peel strength ≥ 3 N/cm by 90° tape test after hard coat cure | Hard coat is applied over AgNW. Poor adhesion causes delamination under thermal cycling. Supplier must confirm compatibility of their AgNW binder with specified hard coat chemistry. |
| Supplier requirement | Tier CAT-B optical coating supplier required; must provide: COA per lot (sheet resistance, transmission, haze), IEC 60068-2-78 qualification data, flex qualification data | Supplier selection must follow NP-PROC-SUP-001 CAT-B criteria. AgNW coating is a novel process vs. ITO — supplier qualification is a one-time process qualification event before production. |
| EC lens bus bar | On EC lens: AgNW layer on outer face connects to EC film inner electrode via silver paste bus bar at lens rim perimeter; bus bar width 2.0 mm; resistance ≤ 5 Ω total path from rim contact pad (F-05) to farthest EC film point | Bus bar resistance adds to EC transition time. ≤ 5 Ω total path resistance keeps transition time within 2 s specification at nominal EC drive voltage. |

BLOCKING REQUIREMENT: AgNW coating supplier must be selected and optical qualification (transmission at 660 nm and 808 nm, haze, flex) completed before any production lenses are ordered. First-article inspection (FAI) for lens must include transmission measurement at 660 nm, 808 nm, and 550 nm on 3 representative lenses from each production lot.

3.5  P-02 — Hard Coat (3–5 µm Silicone)

A hard coat is applied over the AgNW coating on the outer face of both standard and EC lenses, and over the EC film on the inner face of EC lenses. This is standard practice in automotive EC mirrors (ISO 14889 lens coating reference). The hard coat prevents scratch damage to the AgNW network and the active EC layer — neither can be field-repaired if damaged. BOM delta: +$4.00/lens.

Hard coat specification:

| Parameter | Specification | Rationale |
| --- | --- | --- |
| Coating chemistry | Silicone-based UV-cure or thermal-cure hard coat; supplier to confirm chemistry from automotive EC mirror supplier equivalents (e.g., Momentive UVHC series or equivalent) | Silicone hard coat chosen over acrylic: better thermal stability (−40 to +120 °C vs. −20 to +80 °C for acrylic); proven in automotive EC mirror application at 3–5 µm; flexible enough to avoid cracking over 5–10% AgNW flex. |
| Thickness | 3–5 µm dry film thickness (DFT) on outer face; 3–5 µm DFT on EC lens inner face | LOCKED SPEC (CLAUDE.md §7.2). Below 3 µm: insufficient pencil hardness. Above 5 µm: increased internal stress, cracking risk under flex. Standard automotive EC mirror range. |
| Pencil hardness | ≥ 3H per ISO 15184 (pencil hardness test) | 3H equivalent to approximately 200 g/mm² contact hardness — sufficient to resist normal cleaning cloth, key scratch, and ring contact. Vielight-class competitor lenses uncoated (scratched by cotton cloth). |
| Abrasion resistance | ≤ 4% haze increase after 500 cycles Bayer abrasion test (ASTM D4060, Taber CS-10F wheel, 500 g) | Quantitative scratch resistance. 500 Bayer cycles approximates 1 year of daily cleaning. ≤ 4% haze increase = negligible visual impact on stimulation protocols. |
| Adhesion to AgNW | Crosshatch adhesion ≥ 4B per ISO 2409 after 1,000 thermal cycles (−20 to +60 °C, 30 min dwell each) | Thermal cycling between ambient outdoor temperature and active session temperature. ≥ 4B = ≤ 5% area delamination in crosshatch — acceptable for device lifetime. |
| Optical transmission | ≤ 1.5% transmission loss at 550 nm, 660 nm, and 808 nm vs. uncoated substrate | Hard coat anti-reflection treatment (MgF₂ or equivalent) optional if transmission loss exceeds 1.5%. Maintain PBM therapeutic window at 808 nm. |
| Application — outer face | Applied over cured AgNW layer; cure after AgNW adhesion confirmation; inspect for pinholes under 100× before EC lens inner face coating | Pinhole in outer coat on EC lens allows moisture ingress to EC film — catastrophic for EC performance. 100× inspection gate before inner coat application. |
| Application — EC inner face | Applied over EC active layer on inner face of EC lens; thickness 3–5 µm; must not interfere with EC film switching kinetics | Hard coat on inner face protects EC active layer from cleaning damage. Must not add >0.5 V to EC switching voltage — confirm with EC film supplier. |
| UV stability | No visible yellowing (ΔE ≤ 2, CIELAB) after 200 h UVB exposure per ISO 4892-3 | Lens exposed to ambient UV during Mode F outdoor wear. Yellowing attenuates visible wavelengths and degrades visual stimulation protocols. |
| Compatibility check | Hard coat supplier must confirm compatibility with specified AgNW binder and EC film chemistry before production qualification | Three-component stack (PC substrate / AgNW / hard coat on outer; EC film / hard coat on inner) requires system-level adhesion confirmation — not component-level only. |

3.6  P-03 — Inner PDMS Diffuser Film (Plasma-Activated, Eye-Facing)

The inner (eye-facing) surface of both standard and EC lenses carries a PDMS diffuser film with plasma-activated anti-fouling treatment. This is the same PDMS plasma activation process used on the zone module optical windows (NP-HW-FPC-001 §9, PDMS bonding process). The plasma-activated PDMS surface is hydrophilic: it repels sebum and eyelash oils, preventing biofilm and transmission loss on the eye-facing surface. The diffuser layer scatters the 108-LED micro-LED array into a more uniform field across each of the 6 zones per eye.

| Parameter | Specification | Rationale |
| --- | --- | --- |
| Material | PDMS (polydimethylsiloxane), Sylgard 184 or equivalent; 10:1 base:crosslinker ratio; cure 150°C / 15 min | Consistent with zone module PDMS spec (NP-HW-FPC-001 §9.1). PDMS optical transmission >99% at 660–830 nm. Proven anti-fouling surface chemistry after plasma activation. |
| Thickness | 50–100 µm dry film on inner lens face | Sufficient diffusion to smooth LED hot-spots (108 LEDs / lens across 6 zones) while preserving sufficient irradiance for therapeutic dose. Thicker PDMS reduces irradiance by scatter — confirm with optical simulation at 100 µm. |
| Plasma activation | O₂ plasma activation, 100 W, 2 min, 0.2 mbar; perform within 30 min of PDMS cure; apply immediately to lens surface post-activation | O₂ plasma converts PDMS surface from hydrophobic to hydrophilic (contact angle < 15° immediately post-activation). Hydrophilic surface repels sebum (contact angle shift from >100° to <15° vs. native PDMS). Same process as zone module (NP-HW-FPC-001 §9.2). |
| Bond method | Plasma-bonded to PC inner lens surface; SiO₂ interlayer (75 nm RF magnetron sputtered on PC surface) before O₂ plasma, achieves ≥ 150 N/m peel force | Same SiO₂ interlayer bonding process as zone module PDMS (CLAUDE.md §3.1 locked). SiO₂ interlayer on PC achieves same peel force as on PI substrate — confirm on PC representative coupon before production. |
| Anti-fouling durability | Contact angle ≤ 30° after 1,000 eye-proximity cycles (20 cm proximity, 1 s per cycle) | Simulates 6-month daily session use. Contact angle recovery (re-hydrophilisation) expected due to PDMS bulk silanol migration. Confirm with ageing study on representative coupons. |
| Optical haze | ≤ 10% haze at 660 nm and 808 nm; haze must be uniform ± 15% across lens area | Controlled diffusion. Non-uniform haze creates visible hot-spots during EMDR L/R alternation and photic driving protocols. Uniformity verified by scanning transmission measurement on FAI coupons. |
| Compatibility with hard coat | PDMS film on inner face is separate from hard coat on inner face of EC lens; on standard lens, PDMS replaces hard coat on inner face; on EC lens, inner hard coat applied first, PDMS applied over hard coat | EC lens requires inner hard coat for EC film protection; PDMS then applied over hard coat for anti-fouling. Standard lens: no inner hard coat needed; PDMS applied directly to PC substrate via SiO₂ interlayer. |

PDMS bonding qualification: Same 200-cycle IEC 60068-2-14 thermal cycling qualification required as for zone module PDMS (NP-FAI-ZM-001 FAI-TC02). Qualification must be performed on PC substrate with SiO₂ interlayer — zone module qualification (on PI substrate) is not transferable to PC without coupon confirmation.

3.7  F-07 — Lens Rim Guard Anchor Hooks (Goggle Arm)

The lens rim guard (Shore 85A UV-stable TPU, clear, snap-fit over rim profile — see CLAUDE.md §8.3) must be tethered to the headset so it cannot be permanently lost. Tether anchor hooks are molded into the goggle arm at zero incremental tooling cost if specified before first cut. The anchor hook accepts the lens rim guard tether lanyard (1.5 mm braided polyester). Hook position must not interfere with the S1/S2/S3 snap-on shade system — minimum 4 mm clearance from hook to shade body surface confirmed in CAD (MR-19). Tether lanyard length is constrained to ≤ 45 mm free length and must route along the posterior goggle arm surface so that it does not enter the wearer's visual field during rim guard removal (MR-20). Both requirements must be confirmed by CAD before goggle arm tooling is released.

Anchor hook requirements (F-07):

| Parameter | Specification |
| --- | --- |
| Hook geometry | Molded loop, 4 mm internal diameter, 1.5 mm wall, integral to goggle arm outer face at superior-posterior corner |
| Tether break strength | ≥ 50 N (prevents accidental tether breakage under normal use; lens rim guard weight <15 g) |
| Location | 1 per goggle arm (2 per headset); position at superior-posterior corner of arm where arm meets lens rim — accessible without tools |
| Draft angle | ≥ 3° on hook loop interior for mold release without core slide |
| Material | Same as goggle arm body; no additional part or insert required |
| Tether not in scope | Tether lanyard is a separate purchased part (braided polyester, snap-ring terminal); not in tooling spec. Hook geometry must accept Ø 3 mm snap ring. |
| Shade system clearance | Hook position at superior-posterior corner must not contact or deflect any shade body (S1, S2, S3) when shade is attached or detached; minimum 4 mm clearance from hook loop outer edge to nearest shade body surface confirmed in CAD assembly with all three shade body CAD models; clearance must be maintained through full shade snap-on travel, not only in final seated position |
| Tether routing and length constraint | Tether lanyard free length (hook attachment point to lens rim guard snap-ring terminal) ≤ 45 mm when lens rim guard is seated in its rest position on the rim; lanyard routes along the posterior goggle arm outer surface toward the rear of the headset — not across or above the lens zone area; confirm by 3D CAD assembly model and physical mock-up (MR-20) |

4.  MATERIAL SPECIFICATIONS — STRUCTURAL BODIES

| Component | Material | Key Properties Required | Notes |
| --- | --- | --- | --- |
| Lens rim body | Glass-reinforced PA12 (30% GF) or CFRP-composite shell with PA12 GF inner insert | CTE ≤ 40 ppm/°C; Tg ≥ 90°C; impact strength ≥ 60 kJ/m² (ISO 179) | Must maintain ≥ 1 mm polymer wall on all N42 magnet pocket faces across full operating temperature −10 to +45°C. CTE match to PC lens ±20 ppm/°C to prevent rail bind. |
| Goggle arm body | Same material family as lens rim body; CFRP or PA12 GF | Torsional stiffness: ≤ 0.5° arm twist under 5 N lateral lens force | Stiffness prevents pogo pin (F-06) from losing contact with rim pad (F-05) under lateral head acceleration. Torsional stiffness target must be verified in CAD before mold release. |
| Shade body (S1/S2/S3) | S1: opaque PC+ABS; S2: tinted PC; S3 carrier: PC or PA12 | Shade body CTE must match lens rim CTE within ±30 ppm/°C to prevent thermal expansion gap at magnet interface | S3 Rx insert is separate optician-supplied optic; carrier is NeurOne-tooled only. |
| Lens rim guard | Shore 85A UV-stable TPU, clear or tinted (see §8.3) | UV stability: no yellowing ΔE ≤ 2 after 200 h UVB; snap-fit retention force ≥ 5 N, ≤ 20 N extraction | Not in this tooling spec — lens rim guard is a separate tool. Anchor hook (F-07) must be in goggle arm tooling. |

5.  CRITICAL DIMENSIONS

The following dimensions are tolerance-critical and must appear on all lens rim, goggle arm, and shade assembly drawings. Dimensions marked BLOCKING must be confirmed in CAD before mold release.

| Dimension | Nominal | Tolerance | BLOCKING? | Verified by |
| --- | --- | --- | --- | --- |
| Rail groove width at mouth (F-01) | 2.5 mm | ±0.08 mm | Yes | ME CAD + CMM on FAI |
| Rail groove width at base (F-01) | 4.0 mm | ±0.08 mm | Yes | ME CAD + CMM on FAI |
| Rail groove depth (F-01) | 3.0 mm | ±0.10 mm | Yes | ME CAD + CMM on FAI |
| Rail tongue bilateral clearance in groove (F-01/F-02) | 0.08 mm | +0.05/−0.0 mm | Yes | ME CAD assembly simulation |
| Detent engagement depth (F-01) | 1.0 mm | ±0.15 mm | No | FAI detent force measurement |
| N42 magnet pocket diameter (F-03) | 5.05 mm | +0.00/−0.03 mm | Yes | CMM on FAI |
| N42 magnet pocket depth (F-03) | 2.10 mm | +0.00/−0.05 mm | Yes | CMM on FAI |
| Polymer wall — all faces (F-03) | ≥ 1.00 mm | −0.00 mm | BLOCKING | ME CAD cross-section; CMM on FAI all 5 faces |
| EC contact pad diameter (F-05) | 4.0 mm | ±0.10 mm | Yes | CMM on FAI |
| EC contact pad recess depth (F-05) | 0.15 mm | ±0.05 mm | No | CMM on FAI |
| EC pad — nearest magnet pocket separation (F-03/F-05) | ≥ 8 mm | −0.00 mm | Yes | ME CAD measurement |
| EC spring pin bore diameter (F-06) | 1.42 mm | +0.00/−0.02 mm | Yes | CMM on FAI |
| EC spring pin protrusion at mid-stroke (F-06) | 0.80 mm | ±0.10 mm | No | FAI spring force + protrusion measurement |
| Anchor hook internal diameter (F-07) | 4.0 mm | ±0.20 mm | No | Caliper on FAI |
| AgNW hard coat dry film thickness (P-01/P-02) | 3–5 µm | ±0.5 µm | Yes | Profilometer on FAI coupons — 5 locations per lens |
| PDMS inner diffuser thickness (P-03) | 75 µm | ±25 µm | No | Cross-section SEM on FAI coupons |

6.  MOLD DESIGN REVIEW CHECKLIST

This checklist must be completed and signed before steel is cut on any lens assembly tooling package. All items are MANDATORY. Any OPEN item blocks tooling release for the affected package. Reviewer: Mechanical Engineering Lead + Hardware Electrical Engineering Lead + Optical Engineering Lead. Gate: NP-COORD-001 G2 (Pre-Tooling).

The three tooling packages (lens rim, goggle arm, shade assembly) are interdependent. Lens rim and goggle arm reviews must be conducted concurrently. Shade assembly review may follow but must be completed before shade production tooling is released.

| Item | Check | Tooling Package | Reviewer | Status |
| --- | --- | --- | --- | --- |
| MR-01 | Rail groove cross-section (T-slot dimensions, depth, tolerance) confirmed in CAD | Lens rim | ME | OPEN |
| MR-02 | Rail tongue geometry (mating profile, clearance) confirmed in CAD assembly against lens rim groove | Goggle arm | ME | OPEN |
| MR-03 | Rail detent position, engagement depth, and extraction force (≤ 5 N) confirmed by FEA or physical prototype | Lens rim + Goggle arm | ME | OPEN |
| MR-04 | Shade body rail engagement confirmed in CAD assembly for S1, S2, and S3 shade body CAD models | Shade assembly | ME | OPEN |
| MR-05 | N42 magnet pocket ≥ 1 mm polymer wall confirmed on ALL faces in CAD cross-section — BLOCKING | Lens rim | ME | OPEN — BLOCKING |
| MR-06 | N42 magnet pocket 6-position layout confirmed; polarity sequence documented | Lens rim | ME + EE | OPEN |
| MR-07 | Shade mating magnet 6-position layout confirmed against lens rim layout (F-04 mirrors F-03) | Shade assembly | ME | OPEN |
| MR-08 | EC contact pad positions (F-05) confirmed in CAD; ≥ 8 mm separation from nearest magnet pocket verified | Lens rim | ME + EE | OPEN |
| MR-09 | EC spring pin bore positions (F-06) confirmed in CAD assembly against F-05 pad positions; positional tolerance stack-up analysis complete | Goggle arm | ME + EE | OPEN |
| MR-10 | EC contact polarity (positive/negative) marked on lens rim and goggle arm drawings | Lens rim + Goggle arm | EE | OPEN |
| MR-11 | Lens orientation keying confirmed: F-05 pad asymmetry prevents 180° lens rotation (3 independent keying mechanisms) | Lens rim + Goggle arm | ME + EE | OPEN |
| MR-12 | AgNW supplier selected; COA received with sheet resistance, transmission at 550/660/808 nm, and haze — BLOCKING for lens production release | Lens optical stack | Supply Chain + Optical Eng | OPEN — BLOCKING |
| MR-13 | Hard coat supplier confirmed; compatibility with AgNW binder and EC film chemistry confirmed in writing from supplier | Lens optical stack | Supply Chain + Optical Eng | OPEN |
| MR-14 | PDMS SiO₂ interlayer process confirmed on PC substrate coupon (peel force ≥ 150 N/m); separate qualification from zone module PI substrate result | Lens optical stack | Optical Eng + Process Eng | OPEN |
| MR-15 | Anchor hook (F-07) internal diameter 4.0 mm and draft angle ≥ 3° confirmed in CAD; tether lanyard snap-ring compatibility confirmed | Goggle arm | ME | OPEN |
| MR-16 | Full optical stack build sequence documented: substrate → SiO₂ interlayer → AgNW → hard coat (outer face); substrate → EC film → hard coat → PDMS (EC lens inner face); substrate → SiO₂ → PDMS (standard lens inner face) | Lens optical stack | Optical Eng + Process Eng | OPEN |
| MR-17 | PBM transmission confirmed: full optical stack (PC + AgNW + hard coat) transmits ≥ 82% at 660 nm and 808 nm on FAI coupon measurement | Lens optical stack | Optical Eng | OPEN — BLOCKING for Mode F claim |
| MR-18 | EC transition time baseline measured on 5 lenses at production contact resistance (≤ 50 mΩ); transition time ≤ 2 s confirmed | EC lens | EE + Optical Eng | OPEN — EC lens only |
| MR-19 | Anchor hook (F-07) shade system clearance confirmed in CAD assembly: ≥ 4 mm clearance from hook loop outer edge to nearest S1, S2, and S3 shade body surface in fully-seated shade configuration and through full snap-on travel; clearance measurement documented on goggle arm assembly drawing | Goggle arm | ME | OPEN |
| MR-20 | Tether lanyard routing confirmed on goggle arm CAD: lanyard routes along posterior arm outer surface toward rear of headset; free length ≤ 45 mm from hook attachment to lens rim guard snap-ring terminal in rim guard seated position; lanyard path does not cross the lens zone area at any point in rim guard swing arc | Goggle arm | ME | OPEN |

7.  FAI CROSS-REFERENCE

A first article inspection (FAI) checklist for the lens assembly (NP-FAI-LENS-001, to be created) must validate all features in this tooling specification. The following table maps each feature to its FAI validation method. FAI must pass before production lot approval.

| Feature | FAI Validation Method | Acceptance Criterion |
| --- | --- | --- |
| F-01 Sliding rail groove | CMM measurement of groove mouth width, base width, and depth; shade insertion/extraction force measurement (5 units) | All dimensions within tolerance; extraction force ≤ 5 N; detent click confirmed on all 5 units |
| F-02 Rail tongue | CMM measurement of tongue profile; assembled tongue-in-groove clearance measurement | Clearance 0.05–0.10 mm; no binding through full insertion travel |
| F-03 N42 magnet pockets | CMM of pocket diameter and depth; cross-section destructive test (1 unit per lot) to verify ≥ 1 mm wall on all faces | All pocket dimensions within tolerance; wall ≥ 1.00 mm on all 5 faces in destructive cross-section |
| F-04 Shade mating magnets | Shade-to-lens snap force measurement (3 snap cycles per unit, 5 units); shade alignment verification in CAD and on physical assembly | Snap force ≥ 800 g normal per lens; shade alignment within ±0.5 mm of nominal in all 3 axes |
| F-05 EC rim contact pads | CMM pad position; contact resistance measurement at 150 g, 250 g, and 350 g spring force (5 units) | Position within ±0.10 mm; resistance ≤ 50 mΩ at 150 g across all 5 units |
| F-06 EC goggle arm spring pins | Protrusion measurement at mid-stroke; contact force measurement with calibrated gauge; 100-cycle compression test | Protrusion 0.80 ± 0.10 mm; spring force 150–350 g at mid-stroke; resistance ≤ 50 mΩ after 100 cycles |
| P-01 AgNW coating | Transmission measurement at 550, 660, 808 nm (spectrophotometer); sheet resistance measurement (4-point probe, 5 locations); haze (ASTM D1003); flex test 1,000 cycles at R = 50 mm | Transmission ≥ 85%/82%/82% at 550/660/808 nm; sheet resistance 20–30 Ω/sq; haze ≤ 1.5%; resistance after flex within +20% |
| P-02 Hard coat | Profilometer DFT measurement (5 locations); pencil hardness (ISO 15184); adhesion crosshatch (ISO 2409); 200 h UVB exposure (1 unit) | DFT 3–5 µm; pencil hardness ≥ 3H; adhesion ≥ 4B; ΔE ≤ 2 after UV |
| P-03 PDMS inner diffuser | Haze measurement at 660 nm and 808 nm; contact angle measurement before and after 500 proximity cycles; cross-section SEM (1 unit per lot) | Haze ≤ 10%; uniformity ±15%; contact angle ≤ 30° after 500 cycles; thickness 50–100 µm in SEM |
| F-07 Anchor hooks | Dimensional check (caliper); tether lanyard snap ring insertion/extraction test (5 cycles) | Internal diameter 4.0 ± 0.20 mm; snap ring engages and releases without tooling |

8.  OPEN ITEMS

The following items must be resolved before tooling release for the respective package. Items marked BLOCKING prevent lens production from starting.

| OI | Item | Owner | Target Gate | Status |
| --- | --- | --- | --- | --- |
| OI-1 | Select AgNW coating supplier per NP-PROC-SUP-001 CAT-B criteria; obtain qualification data (transmission at 660/808 nm, haze, flex, 85/85 stability) — BLOCKING for lens production release | Supply Chain + Optical Eng | G2 Pre-Tooling | OPEN — BLOCKING |
| OI-2 | Confirm hard coat compatibility with selected AgNW binder and EC film chemistry in writing from hard coat supplier | Supply Chain + Optical Eng | G2 Pre-Tooling | OPEN |
| OI-3 | Perform PDMS SiO₂ interlayer qualification on PC substrate coupon (peel force ≥ 150 N/m, 200-cycle IEC 60068-2-14); separate from zone module PI substrate qualification | Process Eng + Optical Eng | G2 Pre-Tooling | OPEN — BLOCKING for inner PDMS |
| OI-4 | Produce lens rim and goggle arm CAD assembly showing: rail groove + tongue in section, N42 magnet pocket cross-section (all 5 walls ≥ 1 mm), EC contact pad positions, EC spring pin positions, anchor hook geometry | ME | G2 Pre-Tooling | OPEN |
| OI-5 | Confirm shade body rail engagement in CAD assembly for S1, S2, and S3 shade body CAD models; confirm magnet polarity sequence F-03 ↔ F-04 | ME | G2 Pre-Tooling | OPEN |
| OI-6 | Measure full optical stack transmission at 660 nm and 808 nm on first production-representative coupon (PC + SiO₂ + AgNW + hard coat + PDMS); confirm ≥ 82% before Mode F marketing claim is published | Optical Eng | G3 Pre-Production | OPEN |
| OI-7 | Verify EC transition time ≤ 2 s on 5 production lenses with production-representative contacts; log baseline transition time in SHDR factory calibration record | EE + Optical Eng | G3 Pre-Production | OPEN — EC lens only |
| OI-8 | Create NP-FAI-LENS-001 first article inspection checklist for lens assembly; reference this document (NP-TOOL-LENS-001) | ME + QA | G2 Pre-Tooling | OPEN |
| OI-9 | Update NP-COORD-001 engineering coordination checklist to add lens assembly as a tracked deliverable with G2 gate items from this document | Systems Eng | Immediate | CLOSED — G2-13 added to NP-COORD-001 Rev A.3 (2026-05-10). Lens assembly added as G2 tracked deliverable with gate items MR-01 through MR-20 from this document. |
| OI-10 | Confirm anchor hook (F-07) position in goggle arm CAD: (1) verify ≥ 4 mm clearance from hook loop to S1, S2, and S3 shade body surfaces through full snap-on travel (MR-19); (2) confirm tether lanyard routes along posterior arm surface with free length ≤ 45 mm and does not cross lens zone area (MR-20); (3) add clearance and routing dimensions to goggle arm assembly drawing. Satisfies Issue #18 Definition of Done items 1–3. | ME | G2 Pre-Tooling | OPEN |

9.  RISK REGISTER CROSS-REFERENCE

All features in this document address locked design decisions in CLAUDE.md. The following risks from the zone module risk register (NP-FZRM-RISKS-001) are analogous to lens assembly risks and must be tracked in a lens-specific risk entry (to be added to risk register):

| Analogous Zone Module Risk | Lens Assembly Equivalent | Mitigation in This Document |
| --- | --- | --- |
| RISK-04 — PDMS delamination under thermal cycling | PDMS inner diffuser delamination from PC substrate under session thermal cycling | SiO₂ interlayer on PC substrate (P-03); separate 200-cycle qualification from PI substrate result (OI-3) |
| RISK-06 — Photodiode contamination from adjacent components | AgNW bus bar resistance degrading EC transition time — mis-attributed to contact resistance | EC driver logs both contact resistance (from F-05/F-06 measurement) and transition time separately; firmware compares to detect bus bar degradation vs. contact degradation |
| RISK-15 — Module mis-insertion (keying) | EC lens installed rotated 180° (polarity reversal → EC film damage) | Three independent keying mechanisms: F-05 pad asymmetry (MR-11), rail detent position asymmetry, lens rim guard hook asymmetry (F-07) |
| RISK-16 — Gasket IPX4 after repeated swaps | N42 magnet pocket wall integrity after repeated drop events (shade snap-on forces) | ≥ 1 mm wall all faces (F-03, §3.2); 10,000-cycle snap/unsnap qualification; N42 impact tolerance superior to N52 |
| RISK-20 — CFRP Ra confirmation (BLOCKING) | Lens rim slot Ra at rail groove flanks and magnet pocket bore — smooth surface required for shade engagement and magnet adhesive bond | Ra ≤ 0.8 µm at rail groove flanks specified (§3.1); supplier written confirmation required as BLOCKING item equivalent to RISK-20 |
| RISK-23 — Tooling complexity omission | Lens tooling covers 3 interdependent packages (lens rim, goggle arm, shade); omission risk in any package | This document (NP-TOOL-LENS-001) consolidates all 7 features + 3 process specs across all packages; 18-item mold review checklist; concurrent review of lens rim + goggle arm required |

10.  REVISION HISTORY

| Rev | Date | Author | Changes |
| --- | --- | --- | --- |
| A | 2026-05-10 | NeurOne Engineering | Initial release — consolidates 5 locked design changes: F-01/F-02 sliding rail, F-03/F-04 N42 magnet pockets, F-05/F-06 EC driver contacts, P-01 AgNW coating, P-02 hard coat. Adds P-03 PDMS inner diffuser and F-07 anchor hooks. 18-item mold review checklist. Addresses Issue #17. |
| B | 2026-05-10 | NeurOne Engineering | F-07 expanded: shade system clearance requirement added (≥ 4 mm from hook to shade body, MR-19) and tether routing/length constraint added (≤ 45 mm free length, posterior routing, no FOV crossing, MR-20). OI-10 added to track CAD confirmation of both requirements. Mold design review checklist extended from 18 to 20 items. Addresses Issue #18 Definition of Done. |