NP-TOOL-ZM-001

NeurOne Zone Module Tooling Specification

| Document number: | NP-TOOL-ZM-001 |
| --- | --- |
| Revision: | A (Draft — Pre-Tooling) |
| Status: | OPEN — must be signed off before mold design commences |
| Applies to: | Zone Module body mold — all 5 zone positions (ZM-01 through ZM-05) |
| T1 / T2 shared mold: | YES — identical mold for both tiers; differences are firmware-only |
| Date: | 2026-05-06 |
| Related specs: | NP-HW-FPC-001, NP-DRV-SHELL-001, NP-FAI-ZM-001, NP-COORD-001 |

1.  PURPOSE AND SCOPE

This document consolidates all molded and co-molded feature requirements for the NeurOne Zone Module body. Four mitigations (RISK-06, RISK-14, RISK-15, RISK-16) each add a distinct molded feature to the zone module. This document exists because the cumulative complexity creates an omission risk (RISK-23): any feature missed before first-cut requires tooling modification (typically $15,000–40,000 and 6–8 week delay). All features below are MANDATORY and must appear in the mold design review before steel is cut.

Scope: Zone Module body only (thermoplastic shell + co-molded silicone). Not in scope: FPC, PDMS window bonding, LED placement (covered by NP-HW-FPC-001), shell slot geometry (NP-DRV-SHELL-001).

2.  CO-MOLDED FEATURES — MANDATORY CHECKLIST

The following table lists all molded features required in the zone module body. Each feature has a risk traceability reference. The mold design review (Gate: G1 Pre-Layout, item G1-05) must confirm all items before proceeding.

| Feature ID | Feature Description | Risk Ref | Spec Ref | Confirmed in Mold Design? |
| --- | --- | --- | --- | --- |
| F-01 | Mechanical zone key: asymmetric D-section protrusion on insertion edge, unique geometry per zone position (ZM-01 through ZM-05). Prevents physical insertion into wrong shell slot. Must engage shell slot key before FPC tail contacts ZIF. Protrusion height ≥ 1.0 mm; width ≥ 2.0 mm; must clear without binding across 1,000 insertion cycles. | RISK-15 | NP-HW-FPC-001 §11.4 Layer 1 | ☐ |
| F-02 | Raised numeral on module body: ISO-standard raised numeral '1' through '5' (zone position). Numeral height ≥ 0.8 mm; ISO 17049:2013 braille-dot height. Finger-readable by eyes-closed test. Located on thumb-contact surface (posterior face, within 15 mm of extraction point). | RISK-15 | NP-HW-FPC-001 §11.4 Layer 3 | ☐ |
| F-03 | Braille zone abbreviation: ISO 17049:2013 braille cell on same face as numeral (F-02). Zone abbreviations: ZM-01='FL' (Frontal Left), ZM-02='FR' (Frontal Right), ZM-03='CL' (Central Left), ZM-04='CR' (Central Right), ZM-05='OC' (Occipital Center). Braille dot diameter 1.5 mm ± 0.1 mm; dot height 0.48 mm ± 0.05 mm; inter-dot 2.3 mm. | RISK-15 | NP-HW-FPC-001 §11.4 Layer 3 | ☐ |
| F-04 | RISK-14 Option B PD2 aperture: 1.2 mm circular aperture in scalp-facing PDMS window, positioned at X = 33.0 mm, Y = 39.0 mm from module reference corner (lower-left of LED array zone, scalp-facing view). This is the geometric center of the 66 × 78 mm LED array — co-located in XY with PD1 on the LED-facing surface. Aperture tolerance: ±0.2 mm on both axes. Aperture wall: angled 15° inward (funnel entry) to maximise backscattered light collection from scalp tissue. Aperture must be free of flash and burr (Ra ≤ 1.6 µm interior surface). POSITION LOCKED — NP-HW-FPC-001 Rev D §8.4 — 2026-05-09. | RISK-14 | NP-HW-FPC-001 §8.4 | ☐ |
| F-05 | Self-sealing compression gasket (co-molded silicone): Shore 40–50A UV-stable medical-grade silicone (ISO 10993-5), D-section, 2.5 mm wide × 2.0 mm height nominal, continuous bead around full perimeter of zone module insertion edge. Compressed to 1.6 mm (20% compression, 0.4 mm deflection) when module is fully seated in shell slot. No user RTV or field re-sealing action required. Provides IPX4 compliance at module interface. Gasket must maintain seal after 1,000 insertion/removal cycles (per NP-FAI-ZM-001 FAI-IPX-02). | RISK-16 | NP-HW-FPC-001 §12 | ☐ |
| F-06 | Gasket retention groove/undercut: a 0.5 mm undercut groove molded into the zone module body perimeter accommodates the gasket D-section heel. Silicone primer (Dow Corning 1200 or equivalent) applied to groove surface before gasket co-moulding. Combined mechanical retention (groove) + chemical bonding (primer) prevents gasket delamination under 1,000 insertion cycles. Per NP-FAI-ZM-001 FAI-K09: 50-cycle extraction force test — gasket must not detach. | RISK-19 | NP-FAI-ZM-001 §5 | ☐ |
| F-07 | LED baffle wall: internal molded baffle (height ≥ LED stack height + 0.3 mm) separating PD1 position from LED array. Prevents direct LED emission reaching PD1 — PD1 must measure only forward-exiting light through PDMS window. Baffle aperture 2.0 mm × 2.0 mm aligned with PD1 field of view through PDMS. Omission invalidates dose metering system. | RISK-06 | NP-HW-FPC-001 §8.2 | ☐ |
| F-08 | Sliding eject lever (Option A — RISK-22 mitigation): a molded lever arm integrated into the posterior face of the zone module body. Lever is recessed flush with the module surface when closed (does not protrude beyond module envelope). To eject: user hooks one finger under the lever tip and pivots to 30–45°. Lever geometry provides mechanical advantage of ≥3:1, reducing required extraction force from the nominal 2–4 N gasket compression to ≤1 N at the lever tip — within the grip capability of users with moderate Parkinson's or post-stroke hand weakness. Lever dimensions: arm length 10–12 mm; width 8 mm; thickness 2.5 mm; pivot boss diameter 3 mm. Lever tip rounded (R1.5 mm) for finger comfort. Closed latch: a snap-fit detent (PC/ABS, 0.3 mm interference) holds lever flush during insertion and normal wear — prevents accidental ejection from incidental contact. Detent release force: 0.3–0.5 N (finger-actuable, not thumb-actuable at wrong angle). Lever clears ZIF lever mechanism: ME to confirm ≥2 mm clearance between eject lever pivot boss and ZIF back-flip lever in assembled position (OI-7). [DRAWING REQUIREMENT — ZONE MODULE TOOLING] Lever hinge pin boss molded into zone module body. Hinge pin: 3 mm diameter stainless steel (316), press-fit into boss. Lever arm molded separately in same PC/ABS material, assembled post-mold. Detent rib molded on lever underside. | RISK-22 | NP-TOOL-ZM-001 F-08; HFE formative study | ☐ |

3.  MATERIAL SPECIFICATIONS

3.1  Zone Module Body — Thermoplastic Shell

| Property | Requirement | Test Standard |
| --- | --- | --- |
| Base material | PC/ABS alloy or equivalent, UL94-V0 rated | UL 94 |
| Color coding | One solid molded color per zone (ZM-01=Blue, ZM-02=Red, ZM-03=Green, ZM-04=Yellow, ZM-05=Purple). Color must be in bulk material — not surface paint. Saturated enough for color-vision-normal users; tactile features (F-01/F-02/F-03) are primary zone ID for color-impaired users. | ASTM D2244 |
| Surface finish — gasket mating slot rim | Ra ≤ 1.6 µm on shell slot rim contact surfaces (see NP-DRV-SHELL-001 §2.2). Confirm achievable with molded CFRP — tooling manufacturer written confirmation required before G2 gate (RISK-20). | ISO 4287 |
| Thermal stability | Continuous operation ≤ 70°C, IEC 60601-1 skin contact ≤ 43°C at PDMS window external surface | IEC 60601-1 |
| Chemical resistance | Resistant to isopropyl alcohol (70%), saline 0.9%, and sebum simulant. No cracking or crazing after 1,000 wipe cycles. | ISO 10993-5 |
| Co-molded silicone (gasket, F-05) | Shore 40–50A, UV-stable, ISO 10993-5 biocompatible, compression set ≤ 20% after 70h at 70°C (ASTM D395) | ASTM D395, ISO 10993-5 |

4.  CRITICAL DIMENSIONS

The following dimensions are tolerance-critical and must appear on all zone module drawings.

| Dimension | Nominal | Tolerance | Critical for | Drawing callout required? |
| --- | --- | --- | --- | --- |
| Gasket height (uncompressed, F-05) | 2.0 mm | ±0.1 mm | IPX4 seal (RISK-16) | YES |
| Gasket width (D-section, F-05) | 2.5 mm | ±0.15 mm | IPX4 seal | YES |
| Gasket retention groove depth (F-06) | 0.5 mm | +0.0/−0.1 mm | Delamination prevention (RISK-19) | YES |
| Mechanical key protrusion height (F-01) | ≥1.0 mm | Min only | Wrong-zone prevention (RISK-15) | YES |
| PD2 aperture diameter (F-04) | 1.2 mm | ±0.1 mm | PD2 field of view (RISK-14) | YES |
| PD2 aperture center X (F-04) | 33.0 mm | ±0.2 mm | PD2 optical alignment — match FPC contact pad center (RISK-14) | YES |
| PD2 aperture center Y (F-04) | 39.0 mm | ±0.2 mm | PD2 optical alignment — match FPC contact pad center (RISK-14) | YES |
| PD2 aperture center-to-PD2 body offset | ≤0.2 mm | Max only | Optical alignment (RISK-14) | YES |
| Braille dot height (F-03) | 0.48 mm | ±0.05 mm | Blind accessibility (RISK-15) | YES |
| Braille dot diameter (F-03) | 1.5 mm | ±0.1 mm | Blind accessibility (RISK-15) | YES |
| Raised numeral height (F-02) | ≥0.8 mm | Min only | Tactile ID (RISK-15) | YES |
| Eject lever arm length | 10–12 mm | ±0.5 mm | Mechanical advantage ≥3:1 (RISK-22) | YES |
| PD1 baffle height (F-07) | ≥LED stack + 0.3 mm | Min only | Dose metering (RISK-06) | YES |
| PD1 baffle aperture (F-07) | 2.0 × 2.0 mm | ±0.2 mm | Dose metering | YES |

5.  MOLD DESIGN REVIEW CHECKLIST

This checklist must be completed and signed before steel is cut. All items are MANDATORY. Any OPEN item blocks tooling release.

Gate: G1 Pre-Layout (NP-COORD-001 G1-05). Reviewer: ME + EE + Tooling Engineer.

| Item | Requirement | Confirmed? | Notes |
| --- | --- | --- | --- |
| 1 | All 8 features (F-01 through F-08) appear in mold CAD model | ☐ |  |
| 2 | F-01 mechanical key geometry unique per zone position (5 unique keys, reviewed by ME) | ☐ |  |
| 3 | F-02/F-03 numeral + braille finger-readable in prototype sample (3-subject blind test) | ☐ | Prototype validation |
| 4 | F-04 PD2 aperture aligned to PD2 body position per FPC layout (EE confirms FPC layout frozen) | ☐ | F-04 aperture center: X = 33.0 ± 0.2 mm, Y = 39.0 ± 0.2 mm from module reference corner. Confirm in mold drawing. Position locked per NP-HW-FPC-001 Rev D §8.4 (2026-05-09). |
| 5 | F-05/F-06 gasket co-moulding process confirmed with moulder — silicone primer specified | ☐ | Get written process confirmation |
| 6 | F-07 baffle height confirmed vs. LED stack height in BOM (EE confirms LED supplier + package) | ☐ |  |
| 7 | F-08 pull tab does not interfere with ZIF lever access (ME confirms clearance in CAD) | ☐ |  |
| 8 | CFRP shell slot rim Ra ≤ 1.6 µm achievable — written confirmation from tooling manufacturer (RISK-20) | ☐ | BLOCKING if not confirmed |
| 9 | Gasket compression per slot geometry: shell slot width minus module body width = 0.4 mm ± 0.1 mm (ME) | ☐ |  |
| 10 | T1/T2 shared mold confirmed — identical hardware, firmware-only distinction (EE sign-off) | ☐ |  |
| 11 | Zone color in bulk material (not paint) — material spec confirmed with moulding supplier | ☐ |  |
| 12 | All critical dimensions (§4) appear as drawing callouts with tolerances | ☐ |  |

6.  FAI CROSS-REFERENCE

The following first-article inspection items in NP-FAI-ZM-001 directly validate features in this tooling specification. FAI must pass before production lot approval.

| Feature | FAI Item(s) | Pass Criterion |
| --- | --- | --- |
| F-01 (mechanical key) | FAI-A14 | Zero wrong-slot insertions in 20 attempts |
| F-02/F-03 (tactile ID) | FAI-A09, FAI-A10 | Braille readable; dot count correct |
| F-04 (PD2 aperture) | FAI-SY01 (PD2 channel functional) | PD2 ADC reading responds to LED modulation |
| F-05/F-06 (gasket seal) | FAI-IPX-01, FAI-IPX-02, FAI-IPX-03 | IPX4 pass; gasket stays bonded after 10 swap cycles |
| F-07 (baffle) | FAI-A06 (photodiode linearity) | PD output linear with LED current; no direct LED bleed |
| F-08 (pull tab) | HFE formative study (pre-FAI) | Users with Parkinson's/tremor extract module in ≤ 10 s |

7.  OPEN ITEMS

| OI | Description | Owner | Required by | Status |
| --- | --- | --- | --- | --- |
| OI-1 | Written confirmation from tooling manufacturer that CFRP shell slot rim Ra ≤ 1.6 µm is achievable with moulding (RISK-20 — BLOCKING) | ME + Tooling | G2 Pre-Tooling | OPEN |
| OI-2 | Silicone primer specification confirmed with gasket moulding supplier (process compatibility, ISO 10993-5) | ME | G1 Pre-Layout | OPEN |
| OI-3 | FPC layout frozen and PD2 component position confirmed (required before F-04 aperture can be specified in mold) | EE | G1 Pre-Layout | CLOSED — NP-HW-FPC-001 Rev D frozen 2026-05-09. PD2 at X = 33.0 mm, Y = 39.0 mm from module reference corner. See NP-HW-FPC-001 Rev D §8.4 and pin 19 notes. F-04 position handed to mold designer. |
| OI-4 | HFE formative study for pull tab (F-08) — recruit 5 users with Parkinson's or reduced grip strength | UX/Clinical | G2 Pre-Tooling | OPEN |
| OI-5 | Zone color palette confirmed with accessibility review (must pass against shell color under all lighting conditions) | Design | G1 Pre-Layout | OPEN |
| OI-6 | Mold DFM review with tooling house — confirm all features manufacturable with proposed gate locations and draft angles | ME + Tooling | G1 Pre-Layout | OPEN |
| OI-7 | Confirm ≥2 mm clearance between eject lever pivot boss and ZIF back-flip lever in assembled module position. ME to verify in CAD and on prototype. | ME | G1 Pre-Layout | OPEN |

8.  REVISION HISTORY

| Rev | Date | Changes | Author |
| --- | --- | --- | --- |
| A | 2026-05-06 | Initial release. Consolidates RISK-06 (F-07 baffle), RISK-14 Option B (F-04 PD2), RISK-15 (F-01 key, F-02 numeral, F-03 braille), RISK-16 (F-05 gasket), RISK-19 (F-06 retention groove), RISK-22 (F-08 pull tab). Created to mitigate RISK-23 (mold complexity omission risk). | Claude/EE+ME |