SUPERSEDED —

NeurOne

Comprehensive Design Brief

Multi-Modal Closed-Loop Neuromodulation Platform — Consumer + Clinical

Generated: May 2, 2026

# 1. Executive Summary

NeurOne is a two-tier closed-loop neuromodulation platform — the first consumer and clinical brain wearable to integrate photobiomodulation, EEG neurofeedback, transcranial electrical stimulation, vagus nerve stimulation, audio entrainment, and visual stimulation in a single headset with autonomous mobile operation and multi-layer EMF shielding.

## Product tiers

NeurOne Home — Tier 1 (Consumer Wellness)

- FDA-exempt wellness device — no regulatory submission required for launch

- 8 modalities: PBM (660–670nm + 808–830nm), intranasal PBM, EEG (8-ch), brainwave entrainment stimulation, cortical priming stimulation, VNS+HRV, audio, visual

- Pricing: $449 (Core) · $599 (Lite) · $849 (Standard, flagship) · $1,199 (Premium)

- Target: health-conscious adults 35–65, biohackers, wellness practitioners, premium consumer electronics buyers

- Competitive positioning: more modalities, dual wavelength, bilateral intranasal, 8-ch EEG, and no mandatory subscription vs. Neuronic ($3,000+), Vielight ($5,000), sens.ai ($1,500–2,000 + subscription)

NeurOne Pro — Tier 2 (Clinical / Research)

- FDA 510(k) target — 18–36 month pathway using multiple cleared predicates per modality

- Adds: TMS (focal coil, rTMS/TBS), 16-ch qEEG + sLORETA source imaging, 1170nm deep PBM (35–40mm), clinical tACS (≤4mA, 16-ch), cervical VNS option, HIPAA cloud, EHR export (FHIR R4)

- Pricing: $4,999 (Pro Entry) · $13,999 (Pro Full) + $1,800/yr service contract

- Target: neurology clinics (TMS reimbursed CPT 90867–90869), psychiatry, TBI rehabilitation, academic research institutions

## Founding design principles

- Shared platform: both tiers share headset chassis, processors, fit system, app, and USB-C connectivity — one production line, two markets

- Wired-first: USB-C default with zero RF at scalp; BT and Wi-Fi optional via toggle; antennas in control hub not headset

- Autonomous closed-loop: EEG-adaptive stimulation runs in the headset processor without any connected device or phone

- EMF shielding: 5-layer passive + active Helmholtz cancellation — only consumer brain wearable with measured multi-layer shielding ($119 BOM)

- Modular upgradeability: snap-in zone modules and magnetic accessory ports enable field upgrades without chassis replacement

- No mandatory subscription: full function on-device from day one — direct competitive counter to sens.ai model

# 2. Modality Stack — Full Specification

All modalities run simultaneously or in any user-defined sequence. The closed-loop engine on the headset processor can adaptively control any modality parameter based on live EEG state, with <200ms end-to-end latency.

| Modality | Consumer name | Tier | Tier 1 specification | Tier 2 additions / differences |
| --- | --- | --- | --- | --- |
| PBM transcranial | Photobiomodulation | T1+T2 | 660–670nm + 808–830nm LEDs · 5 independent zones · 10–200 mW/cm² closed-loop · real-time J/cm² dose metering · 0–10,000Hz pulse · photodiode feedback per zone | Adds 1170nm laser diode deep channel · 35–40mm subcortical depth · TEC stabilisation · 10–1,000 mW/cm² · 5 full zones at all 3 wavelengths |
| PBM intranasal | Intranasal light therapy | T1+T2 | Bilateral Y-probe · 660nm + 808–830nm per probe · 15/20/25mm depth stops · photodiode dose sensing · optical code + pogo sleeve auth · hygiene sleeve consumable (30-pack) | Higher irradiance ceiling (150 mW/cm²) · EEG-adaptive theta protocol · formal clinical integration into session management |
| EEG / qEEG | Brain monitoring | T1+T2 | 8-ch semi-dry hydrogel electrodes · Fp1/2, F3/4, C3/4, P3/4 · 500Hz · delta→high gamma · real-time neurofeedback · power bands + coherence | 16-ch wet gel upgrade · <1µV RMS noise · CMRR >120dB · sLORETA source imaging · EDF+ export · ERP paradigms · LSL streaming · research scripting API |
| tACS | Brainwave entrainment stimulation | T1+T2 | 0.5–40Hz · ≤1mA · 3 electrode pairs · charge-balanced biphasic · real-time impedance monitoring · consumer parameter ranges | 0.5–100Hz · ≤4mA · 16-ch arbitrary waveform · clinician-controlled full parameter exposure |
| tDCS | Cortical priming stimulation | T1+T2 | 0.1–2mA DC · 5–30 min sessions · 30s ramp up/down · ≤3 electrode pairs · 40µC/cm² limit hardware-enforced | 0.1–4mA · up to 8 simultaneous pairs · 10–60s programmable ramp · HD-tDCS · full montage library |
| VNS + HRV | Vagus nerve stimulation | T1+T2 | Auricular clip · 1–25Hz · ≤2mA · wellness params · PPG HRV in same clip · adaptive breathing guide · force contact confirmation | 0–5mA · clinical biphasic waveforms · optional cervical electrode accessory · full autonomic monitoring suite |
| Audio entrainment | Neural audio entrainment | T1+T2 | Over-ear planar magnetic 40mm · bone conduction at mastoid · binaural beats · isochronic tones · pink/brown noise · VNS clip cutout in cup rim | EEG-adaptive bilateral audio · clinical EMDR bilateral alternation mode · session guidance overlay |
| Visual stimulation | Visual neural entrainment | T1+T2 | 108 micro-LEDs per lens (660nm + 808–830nm) · 6 zones/eye · standard 35% VLT or EC 5–75% VLT · photic driving 0.5–100Hz · EMDR L/R alternation · retinal PBM · IR eye-open detection · IEC 62471 hardware limits | Clinical EMDR protocols · EEG-adaptive photic phase-locked to live oscillation · photoparoxysmal EEG detection → auto-halt <200ms |
| TMS | — (clinical only) | T2 only | Not available in Tier 1 | Focal figure-8 coil · 0.1–0.5T · rTMS + TBS protocols · active coil cooling · capacitor bank in separate hub · non-conductive CFRP window at coil site · TMS-gated EMF cancellation |

## 2.1 PBM Frequency Preset Library

| Preset | Frequency | Primary zones | Secondary modalities | Evidence + rationale |
| --- | --- | --- | --- | --- |
| Gamma clarity | 40 Hz | Prefrontal + temporal | BES gamma · audio 40Hz | MIT Tsai: 40Hz reduces amyloid-β and tau in Alzheimer's models. Vielight primary clinical frequency. Strongest transcranial PBM evidence base. Gamma entrains memory binding and cognitive integration. |
| Alpha calm | 10 Hz | Parietal + occipital | VNS · HRV coherence | Most-published transcranial PBM frequency. Drives thalamocortical alpha loops. Measurably increases EEG alpha power post-session. Entire Vielight commercial track record uses 10Hz. |
| Theta memory | 6 Hz | Temporal + intranasal | Intranasal bilateral · theta BES | Theta is hippocampal dominant rhythm for memory encoding. Intranasal delivers theta-coupled light directly to olfactory bulb → hippocampus pathway. Age-related cognitive decline protocol. |
| Sleep deep | 2 Hz | Full cap + intranasal | Delta audio · VNS parasympathetic | Delta supports slow-wave sleep and glymphatic waste clearance. Best pre-sleep protocol. Glymphatic function increasingly linked to dementia prevention. |
| Gamma+theta coupled | 40+6 Hz split | Transcranial/intranasal | Audio theta-gamma coupling | Mirrors hippocampal theta-gamma cross-frequency coupling. 40Hz gamma nested in 6Hz theta phase matches endogenous architecture. No existing device implements this protocol. |
| Focus prime | 20 Hz | Prefrontal bilateral | tDCS cortical priming | Beta drives active cognitive engagement. tDCS anodal priming amplifies effect. Pre-task executive function enhancement. |
| Vascular baseline | CW (0 Hz) | Intranasal only | None | Continuous wave maximises nasal capillary irradiation for systemic PBM effect. Research baseline condition. |

# 3. Hardware Specifications

| System | Specification | Key rationale |
| --- | --- | --- |
| Main processor | NXP i.MX RT1062 · Cortex-M7 · 600MHz · FPU+DSP+SIMD · 1MB SRAM + 32MB LPSDR4 SDRAM · 8GB eMMC 5.1 · USB-HS OTG (480Mbps) · FreeRTOS · ~1.1% CPU at full load | 98.9% headroom reserved for future on-device ML. Rich peripheral set (FlexPWM, multi-SPI) replaces external ICs. Autonomous closed-loop: FFT + state classification + parameter update at 10Hz with no connected device. |
| Safety MCU | STM32G071 · Cortex-M0+ · 64MHz · 36KB SRAM · 128KB flash · bare metal · owns all stimulation GPIO enables · 1.5s watchdog · Q15 fixed-point CMSIS-DSP | Upgraded from G031 (8KB SRAM error discovered in analysis — G071 is drop-in at +$0.45). IEC 62304 Class C scope limited to ~500 lines. Hardware key: app crash cannot cause unsafe stimulation. Cutoff <50ms from any fault. |
| Fit system | Boa-style occipital dial (10cm range, 0.5mm/click, 50,000 cycle rated) · 5-position forehead bridge · spring-loaded electrode pods (80–120g, ±12mm travel) · integrated Shore 30A silicone grip · temporal stability wings | 1 adult SKU covers 52–62cm (90%+ of adults). Spring-decoupled electrode pressure: optimal EEG contact independent of overall tightness. Solves fundamental comfort-vs-signal conflict. No velcro, no removable pads, autoclavable arch. |
| EMF shielding | CFRP outer shell (30–50dB RF) + mu-metal liner (15–25dB ELF) + silver fabric inner liner (40–60dB RF) + EMI absorber foam + port filters (30–50dB) + active Helmholtz cancellation (adds 20–25dB ELF). Combined: 35–45dB ELF magnetic, 40–60dB RF. +$119 BOM. | Only consumer brain wearable with measured multi-layer EMF shielding. All shell layers bonded to EEG DRL output (standard shielded-EEG practice). Non-conductive CFRP window at TMS site prevents eddy currents. Three firmware fixes: TMS-gated cancellation, adaptive notch, Helmholtz subtraction. |
| Connectivity | USB-C 3.2 Gen1 default (zero RF, <1ms latency, PD charging) · Bluetooth 5.3 LE Audio (Mode 2) · Wi-Fi 6 2.4/5GHz (Mode 3) · Single rear toggle button · LED ring per mode · Antennas in control hub (not headset) | Hub-mounted antennas reduce scalp RF 20–30dB vs headset-mounted even in wireless mode. Phone-as-power: 1.5m cable from phone-in-pocket reduces cellular RF to ambient background (<0.05µV/cm²). MagSafe breakaway at shoulder prevents headset yanking. |
| Autonomous operation | Pre-program via USB-C (<5s, 50–500KB protocol) → disconnect → run on any USB-C PD power bank. Closed-loop EEG-adaptive stimulation runs in headset. Data logged to 8GB eMMC. Download on reconnect (EDF+ + parameter logs). | No competitor offers autonomous closed-loop. Pre-program-and-disconnect is unique in consumer neuromodulation. 10,000mAh PD bank: ~90 min standard protocol runtime. Processor retained at $85 BOM — removing it saves 10% COGS while permanently destroying the primary competitive moat. |
| Status indicators | Left temple: green power LED · Right temple: amber in-use LED · Amber pulse rate mirrors session frequency · Power LED doubles as fault (red blink) · Stealth mode (app-controlled) · Both: forward-facing, visible at eye level of standing observer | Clinical rationale: caregiver across room confirms session is active and at correct frequency without touching app. Spatially distinct LEDs = additive reading, no learned color code. Amber frequency confirmation is unique in the industry. |
| Visual lens system | Standard: 35% VLT polycarbonate shield · 108 micro-LEDs/lens (660+810nm) · 6 zones/eye · PDMS diffuser · ITO EMF coating · S1 opaque shade included. EC premium (+$89): bistable electrochromic 5–75% VLT · 2s transition · ~15mW hold. | Micro-LED not a display — a controllable light field. Avoids waveguide cost ($180–350/lens) with no functional loss. Mode F (NIR retinal walk): invisible daily-use protocol, 808–830nm imperceptible to user. Bistable EC: holds state without power, always clears to 75% on power restore (safety). |

# 4. Configurations + Pricing

## 4.1 Integrated system configurations

| Configuration | BOM est. | Mfg. COGS | Retail price | Gross margin | Included modalities | Target customer |
| --- | --- | --- | --- | --- | --- | --- |
| Core — EEG only | $165 | $255 | $449 | 43% | 4-ch EEG · USB-C · BT · Wi-Fi · EMF shielding · processors · storage · indicators | Cautious entry buyer; neurofeedback researchers |
| Home Lite | $260 | $365 | $599 | 39% | All Core + PBM 5 zones (660+810nm) · 8-ch EEG · VNS + HRV | Wellness-focused consumer; Neuronic/Vielight alternative seeker |
| Home Standard ★ | $395 | $530 | $849 | 38% | All T1 modalities: PBM · intranasal · 8-ch EEG · BES · tDCS · VNS+HRV · audio · visual · S1 shade | Primary flagship; biohacker; health-conscious adult 35–65 |
| Home Premium | $450 | $610 | $1,199 | 49% | All T1 modalities + EC lens upgrade + extended 2-yr warranty + priority support | High-income buyer; premium lens experience preferred |
| Pro Entry | $820 | $1,350 | $4,999 | 73% | All T1 + 16-ch qEEG · 1170nm deep PBM · clinical tACS · HIPAA cloud · EDF+ export | Neurology clinic; concussion rehab center; research PI |
| Pro Full | $1,480 | $2,600 | $13,999 | 81% | All T2: Pro Entry + TMS hub · multi-patient dashboard · research mode · LSL · scripting API · FHIR R4 EHR export | Academic medical center; psychiatry clinic; neuromodulation research lab |

## 4.2 Power requirements by operating mode

PBM is 85–90% of total power draw in every active mode. All other modalities combined account for only 2–5W. The standard 45W MacBook charger (15V/3A) is the recommended everyday supply for Tier 1 users.

| Operating mode | Total draw | Min USB-C PD | Runtime (10,000mAh) | Notes |
| --- | --- | --- | --- | --- |
| Standby / idle | 1W | 5V/0.5A | ~330 hours | Any USB source. Phone USB-A adapter sufficient. |
| EEG monitoring only | 2.5W | 5V/1A | ~130 hours | Laptop USB port sufficient. Closed-loop neurofeedback fully functional. |
| Standard T1 protocol ★ | 20W | 15V/2A (45W) | ~90 min | 5 zones 660+810nm + all modalities. 45W MacBook charger recommended. |
| All modalities T1 peak | 50W | 20V/3A (65W) | ~36 min | Maximum irradiance all zones simultaneously. |
| T2 standard (+ 1170nm) | 46W | 20V/3A (65W) | ~39 min | Typical clinical session with deep PBM. |
| T2 peak (no TMS) | 74W | 20V/5A (100W EPR) | ~24 min | All 3 wavelengths at maximum. |
| TMS hub (Tier 2 add-on) | 30–50W | Dedicated supply | Wall recommended | Separate USB-C 100W or barrel-jack. Not from headset PD bus. |

## 4.3 Modular pricing (build-your-own)

Integrated systems are the primary sales motion and offer 36–83% savings vs. modular assembly. Modular pricing exists for field upgrades, clinical procurement line-item accounting, and replacement components.

| Module | BOM est. | Modular retail | Integration value | Notes |
| --- | --- | --- | --- | --- |
| PBM zones — 5-zone 660+810nm | $140 | $299 | $220 | Largest single BOM item after core. Drives 85% of session power draw. |
| EEG 8-channel expansion | $45 | $129 | $90 | Semi-dry hydrogel pods. Same snap-in zone slots. |
| Brainwave entrainment + cortical priming (BES+tDCS) | $35 | $149 | $100 | Shared electrode array with EEG via fast analog switch matrix. |
| VNS auricular clip + HRV | $22 | $89 | $60 | Magnetic 6-pin accessory port. PPG + stimulation in single unit. |
| Intranasal PBM applicator | $35 | $149 | $90 | Sleeve consumable subscription drives recurring revenue. |
| Over-ear audio cups | $55 | $179 | $130 | Bayonet quarter-turn. Clearance for VNS clip built into lower rim. |
| Visual goggle shield (standard lens) | $38 | $129 | $90 | 108 micro-LEDs/lens. ITO EMF coating. S1 shade included. |
| EC lens upgrade (premium) | $47–67 | $89 upgrade | $130 value | Bistable. Replaces S2 polarising clip need. |
| EEG 16-ch pro expansion (Tier 2) | $65 | $249 | $175 | Wet gel. <1µV RMS. Tier 2 only. |
| 1170nm deep PBM — all 5 zones (Tier 2) | $250 | $899 | $700 | Laser diode + TEC. Subcortical depth 35–40mm. |
| TMS hub (Tier 2, separate unit) | $200 | $2,999 | $2,400 | Separate chassis. Figure-8 coil. Connects via USB-C to headset. |

# 5. Accessory Catalog

| Accessory | Price | Compatible | Key specification | Market rationale |
| --- | --- | --- | --- | --- |
| Intranasal PBM applicator | $149 + $19/mo sleeves | T1 + T2 | Bilateral Y-probe · 660nm + 808–830nm per probe · 15/20/25mm depth stops · photodiode contact sensing · optical sleeve auth (no NFC) · 30-pack sleeves included | Olfactory-hippocampal pathway inaccessible transcranially. Exceeds Vielight's single fixed-depth intranasal with bilateral design, dual wavelength, 3 depth stops. Sleeve subscription = high-margin recurring revenue (68–79% gross margin on consumables). |
| VNS auricular clip + HRV | Included T1 Std+ | T1 + T2 | Auricular electrical VNS · PPG HRV · 808–830nm optical · force contact sensor · PDMS hydrogel pad · same magnetic 6-pin port as intranasal | VNS current warms earlobe → improves PPG quality. Real-time closed-loop vagal-cardiac feedback. HeartMath earlobe HRV validation. Electrical auricular tVNS has more evidence than Vielight's optical VNS alternative. |
| Over-ear audio cups | Included T1 Std+ | T1 + T2 | 40mm planar magnetic · bone conduction at mastoid · VNS clip clearance channel in lower rim · binaural beats + isochronic tones · EEG-adaptive frequency | Binaural beats require stereo isolation. Bone conduction adds dual-pathway auditory entrainment. MIT: combined audio+visual 40Hz extends amyloid reduction beyond visual cortex to hippocampus. |
| EC lens upgrade | +$89 / $129 standalone | T1 + T2 | Bistable electrochromic film · 5–75% VLT · app-controlled · 2s transition · ~15mW hold power · clears to 75% on power restore (safety failsafe) | Eliminates S2 polarising clip need. App-controlled opacity removes physical shade management. Bistable = negligible power impact (~0.45Wh/30-min session). Premium lens experience for high-end users. |
| S1 opaque shade | Included / $18 replacement | Std + EC lens | <0.5% VLT · magnetic snap (6× N52 magnets in lens rim) · foam light seal · protective case function when stored | Instant full immersion (EC transition 2s too slow for some protocols). Retained as EC failsafe. Maximum stimulation fidelity for seated clinical sessions. |
| S2 polarising clip | $24 | Std lens only | ~12% VLT · polarised polycarbonate · magnetic snap · outdoor photic use | Outdoor photic driving: sunlight (50,000–100,000 lux) overwhelms LED at 35% VLT. S2 enables outdoor mobile photic protocols. Not needed with EC lens. |
| S3 prescription clip | $49 carrier + $49–139 Rx | Std + EC lens | 1.67 high-index corrective lens · 14mm vertex distance · single-vision or progressive · app-based Rx submission · partner optician network · 5–7 day fulfilment | 75% of adults 40+ need vision correction. Without S3, majority of target demographic excluded from open-eye sessions. Rx renews every 12–24 months = recurring revenue at zero marginal marketing cost. |
| Cervical VNS electrode | $49 | T2 only | Over carotid triangle CN X trunk · gel contacts · same magnetic connector as auricular clip · swappable head | Clinical VNS protocols sometimes require cervical delivery. gammaCore cleared predicates support regulatory path. Auricular is standard; cervical is clinical accessory. |
| Autoclavable arch set | $39/pair | T2 clinical | Shore 30A silicone forehead + occipital arches · quarter-turn field-replaceable · autoclave-rated | JCAHO and IPC infection control: Neuronic velcro cannot be autoclaved. Two-arch rotation enables back-to-back patient sessions in clinical setting. |
| Hygiene sleeve 30-pack | $19/pack or $19/mo | T1 + T2 | Medical-grade silicone · laser-etched optical auth code · conductive pogo collar · single-use · zero NFC/EMF | Only consumable in ecosystem. 68–79% gross margin. Clinical settings require documented single-use. Optical+resistive auth: zero EMF, faster than NFC, lower BOM per sleeve. |

# 6. Competitive Positioning

| Feature | NeurOne Home | NeurOne Pro | Vielight Neuro Pro 2 | Neuronic 1070 | Sens.ai |
| --- | --- | --- | --- | --- | --- |
| Price | $849 | $4,999–13,999 | ~$5,000 | $3,000–4,995 | ~$1,500–2,000 |
| PBM wavelengths | 660+810nm (2λ) | 660+810+1170nm (3λ) | 810nm only (1λ) | 1070nm only (1λ) | ~810nm only (1λ) |
| PBM zones | 5 zones, ~150 LEDs | 5 zones + deep LD | 10 transcranial + 2 intranasal | 4 quadrants, 256–300 LEDs | 7 LEDs, midline only |
| Intranasal PBM | Bilateral, 2λ, 3 depths | Bilateral, 2λ, 3 depths | Yes (single fixed depth) | No | No |
| Dose tracking | Real-time J/cm² per zone | Real-time J/cm² per zone | No | No | No |
| EEG channels | 8 semi-dry | 16 wet gel + qEEG | None (roadmap) | None | 3 dry midline |
| Closed-loop EEG | Yes — autonomous | Yes — all modalities | No (roadmap only) | No | EEG → PBM freq only |
| tACS / BES | Yes (consumer name) | Yes, clinical grade | No | No | No |
| tDCS | Yes | Yes | No | No | No |
| TMS | No | Yes | No | No | No |
| VNS + HRV | Auricular, electrical | Auricular + cervical | Optical PBM VNS | No | HRV only, no VNS |
| Audio entrainment | Binaural + bone conduction | + clinical EMDR | No | No | No |
| Visual / photic | 108 LEDs/lens + EMDR | + EEG-adaptive | No | No | No |
| EMF shielding | 5-layer + active | 5-layer + active | None | None | None |
| Autonomous mobile | Pre-program + power bank | Pre-program + power bank | BT only (app required) | Portable, BT only | BT only |
| Wired-first (USB-C) | Yes — zero head RF | Yes — zero head RF | Wireless only | Power only, BT data | Wireless only |
| Prescription support | S3 clip (both lenses) | S3 clip (both lenses) | No | No | No |
| No mandatory subscription | Yes — full features | Yes — full features | Yes | Paywall on PLUS | Required $99–199/yr |
| Published RCTs | None (new product) | None (new product) | 35+ trials | Limited | Ongoing |

## 6.1 Key competitive advantages

- Only device with closed-loop EEG-adaptive stimulation across all modalities — Vielight is developing this but does not yet have EEG hardware

- Only device with dual-wavelength (660+810nm) transcranial PBM — 2024 rat study confirms broader hippocampal activation vs. single wavelength

- Only device with bilateral intranasal PBM at multiple depth stops — Vielight's fixed-depth single probe cannot match anatomical targeting

- Only device with BES (tACS), tDCS, VNS, audio, and visual stimulation combined with EEG — no competitor has more than 2 of these

- Only brain wearable with measured multi-layer EMF shielding including active Helmholtz cancellation

- Only device with fully autonomous closed-loop mobile operation via any USB-C power bank, no phone required

- Prescription clip accessory (S3) — first neuromodulation headset to solve the vision correction exclusion problem

## 6.2 Key competitive gaps (at launch)

- Zero published clinical trials — Vielight has 35+ published RCTs; this is the most significant gap and requires 2–3 years of funded research to address

- Vielight's 400 mW/cm² surface irradiance exceeds NeurOne Tier 1's 200 mW/cm² — Tier 2 at 1,000 mW/cm² exceeds it but at higher price tier

- Neuronic LED count (256–300) exceeds NeurOne Tier 1 (~150) at single wavelength — NeurOne's advantage is depth targeting and dual wavelength, not raw emitter count

# 7. Regulatory + Go-to-Market Strategy

## 7.1 Tier 1 — FDA-exempt wellness pathway

- Positioning: general wellness device under FDA enforcement discretion (same as Muse, Dreem, sens.ai)

- Constraint: no medical claims — 'supports focus and relaxation' not 'treats depression'

- tACS marketed as 'brainwave entrainment stimulation' — accurate description, avoids regulated clinical terminology

- 1170nm laser diode absent from Tier 1 — removes Class 3B laser safety classification requirement

- Time-to-market: 12–18 months from design freeze

## 7.2 Tier 2 — FDA 510(k) pathway

- Strategy: modular predicate device filing — each high-power modality against separate cleared predicate

- TMS predicates: NeuroStar (Neuronetics), BrainsWay (deep TMS) — extensive cleared predicate base

- tACS predicates: Soterix Medical, Neuroelectrics Starstim — cleared tDCS/tES predicates

- VNS predicates: electroCore gammaCore — cleared transcutaneous cervical VNS

- Timeline: 18–36 months post-Tier 1 launch. Budget: $2–5M including required clinical data

- Clinical CPT reimbursement: TMS sessions billable under CPT 90867–90869 — existing payer coverage

## 7.3 Go-to-market sequencing

- Year 1–2: Tier 1 consumer launch — build brand, revenue, and user base. Target: sens.ai, Neuronic, Vielight customers

- Year 2–3: Seed Tier 2 units into research institutions under IRB exemptions — generate published clinical data

- Year 3–4: Tier 2 commercial launch post-510(k). Sales channels: neurology, psychiatry, TBI rehab, academic research

- Recurring revenue: sleeve subscriptions + prescription clip renewals + Tier 2 service contracts ($1,800/yr) + optional advanced protocol library ($9.99/mo)

# 8. Safety Architecture Summary

## 8.1 Dual-processor safety isolation

- STM32G071 safety MCU runs bare-metal, completely independent of main processor — owns all stimulation GPIO enable lines in hardware

- Main processor sends SPI heartbeat every 200ms — 3 missed heartbeats (1.5s timeout) triggers hardware cutoff of all stimulation within 50ms

- IEC 62304 Class C scope: ~500 lines of bare-metal safety firmware — formally verifiable, not dependent on OS or application code

## 8.2 Modality-specific safety interlocks

- EEG: photoparoxysmal EEG response detection → goggle LED cutoff within 200ms (before clinical seizure event)

- tDCS: 40µC/cm² charge density enforced in hardware — app cannot exceed this regardless of software state

- Visual: IEC 62471 retinal MPE hardware limit at 50% of exempt group threshold — physically impossible to exceed from app. IR eye sensor auto-reduces 20× if open eye detected.

- TMS: active EMF cancellation gated off 5ms before each pulse — prevents fluxgate saturation and spurious cancellation field

- PBM: NTC thermistor per zone — hardware current throttle at 62°C junction, enforces IEC 60601 42°C scalp limit

- VNS: impedance checked before each session — low impedance hold if electrodes not in contact

## 8.3 EMF shielding interference management

- Non-conductive CFRP window at TMS coil site prevents eddy current braking (20–40% field loss avoided)

- All conductive shell layers bonded to EEG driven-right-leg (DRL) output — eliminates capacitive common-mode coupling

- Adaptive notch filter at BES/tACS stimulus frequency prevents active cancellation from cancelling the therapeutic stimulus

- Synchronous Helmholtz artefact subtraction from EEG: removes coupling artefact without notch filter, preserving genuine 50/60Hz brain signal

# Document status

This design brief reflects all decisions made through the full design session. The following items remain open for decision:

- Product family name ("NeurOne" is a working name — branding decision required before any external conversations)

- Scientific advisory board composition and clinical trial funding plan for Tier 2 evidence generation

- Manufacturing partner selection and PCB layout vendor (priority: STM32G071 chip change before first layout)

- Partner optician network for S3 prescription clip program

- MFi certification timeline (optional — required only if iOS background execution needed for phone-tethered use case)