NeurOne

Zone Module FPC Assembly
First Article Inspection (FAI) Checklist

NP-FAI-ZM-001  Rev A  |  2026-05-06  |  Pre-Production Draft
References: NP-HW-FPC-001 Rev C · NP-PROC-FPC-001 Rev A · NP-DRV-SHELL-001 Rev A
Risk mitigations consolidated: RISK-01, 02, 04, 05, 06, 07, 08, 10, 11, 12

| Build lot / serial: |  | FAI date: |  |
| --- | --- | --- | --- |
| Headset serial number: |  | Inspector name: |  |
| FPC lot number: |  | Reviewed by: |  |
| Shell revision: |  | Final result: | PASS / FAIL |

SCOPE: This checklist covers the first article of each zone module FPC assembly. ALL items must pass before production release. Any FAIL triggers a disposition review — no waiver without Engineering and Quality sign-off. Items marked [BLOCKING] must pass before any other FAI items proceed.

SECTION 1 — PRE-INSPECTION DOCUMENT CHECK

Verify all required documents and BOM revisions are current before physical inspection.

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-D01 | FPC specification NP-HW-FPC-001 Rev C on file | Document check | Rev C or later confirmed |  |  |
| FAI-D02 | Procurement requirements NP-PROC-FPC-001 Rev A on file | Document check | Rev A or later confirmed |  |  |
| FAI-D03 | Shell tooling design review NP-DRV-SHELL-001 Rev A fully signed off (all 8 signatories) | Document check | All sign-off boxes in §8 completed |  |  |
| FAI-D04 | BOM confirms BCR421W (not BCR421U) — verify against approved BOM revision | BOM review | BCR421W in all LED driver positions |  |  |
| FAI-D05 | BOM confirms Hirose FH34S-20S-0.5SH(50) — no Molex SlimStack or 0.35 mm pitch connectors | BOM review | FH34S or JAE FF03 only in connector positions |  |  |
| FAI-D06 | FPC fab drawing specifies RA copper (IPC-4204/11 Type I) — check drawing fab notes | Drawing review | 'ROLLED ANNEALED' or 'RA per IPC-4204/11' visible as fab note 1 |  |  |
| FAI-D07 | RA copper material certificate accompanies FPC lot | Certificate check | Cert references lot number and IPC-4204/11 Type I |  |  |
| FAI-D08 | LED Vf bin distribution certificate accompanies LED lot (660 nm and 808–830 nm) | Certificate check | Cert present, lot mean ± range within ±0.10 V spec |  |  |

SECTION 2 — INCOMING COMPONENT INSPECTION

Verify purchased components before assembly. Failures here prevent assembly from proceeding.

2a. LED Emitters (660 nm and 808–830 nm)  [RISK-08 Vf binning · RISK-10 RA copper ref]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-L01 | 660 nm LED: measure Vf on 10-unit sample at 150 mA / 10 ms pulse / 25°C | Benchtop measurement | All 10 within ±0.10 V of lot mean on certificate |  |  |
| FAI-L02 | 808 nm LED: measure Vf on 10-unit sample at 150 mA / 10 ms pulse / 25°C | Benchtop measurement | All 10 within ±0.10 V of lot mean on certificate |  |  |
| FAI-L03 | 660 nm peak wavelength: spectrometer check on 5-unit sample | Spectrometer | λ_peak 655–670 nm, within ±5 nm of lot certificate |  |  |
| FAI-L04 | 808 nm peak wavelength: spectrometer check on 5-unit sample | Spectrometer | λ_peak 805–835 nm, within ±10 nm of lot certificate |  |  |
| FAI-L05 | Reel label part number and date code check | Visual | Part number matches BOM; floor life not expired (MSL 3) |  |  |

2b. BCR421W LED Driver IC  [RISK-07 variant confusion]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-C01 | [BLOCKING] Part marking: confirm 'W' suffix on every sampled IC | Visual / magnifier (10×) | Marking reads BCR421W. Any BCR421U → reject entire lot immediately |  |  |
| FAI-C02 | Package type: SOT-323 (SC-70) 3-pin | Visual / caliper | SOT-323 body size 2.0 × 1.25 mm confirmed |  |  |
| FAI-C03 | Authorised distributor CoC present | Certificate check | Distributor is Mouser, DigiKey, or Avnet; CoC on file |  |  |

2c. Hirose FH34S Connector  [RISK-01 connector family]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-K01 | [BLOCKING] Part number on reel label: confirm FH34S-20S-0.5SH(50) | Visual | Exact part number confirmed. Any Molex or 0.35 mm connector → reject |  |  |
| FAI-K02 | Position count: 20 contacts | Visual / pin count | Count = 20 |  |  |
| FAI-K03 | Lever actuation: back-flip lever opens and closes smoothly | Manual test, 3 samples | No stiffness, no binding, lever fully releases at 180° |  |  |
| FAI-K04 | Contact resistance: insert FPC tail, measure all 20 contacts | 4-wire measurement | All contacts < 50 mΩ after 5 insertion cycles |  |  |

2d. FPC Laminate  [RISK-10 RA copper · RISK-02 power trace]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-F01 | [BLOCKING] RA copper material certificate: confirm 'Rolled Annealed' per IPC-4204/11 Type I | Certificate check | Certificate explicitly states RA. 'ED' certificate → reject lot |  |  |
| FAI-F02 | Copper weight: 1 oz (35 µm ± 10%) confirmed on certificate | Certificate check | 35 µm ± 3.5 µm |  |  |
| FAI-F03 | Connector tail hard gold thickness: XRF measurement on 2 samples | XRF (2 pads/unit) | ≥ 0.5 µm gold. Any pad < 0.4 µm → quarantine lot |  |  |
| FAI-F04 | IPC class 2 certificate accompanies lot | Certificate check | Certificate references IPC-A-600 Class 2 or better |  |  |

SECTION 3 — ASSEMBLED ZONE MODULE INSPECTION

Inspect first-article assembled zone module PCB before installation into shell.

3a. Power Trace and Thermal Verification  [RISK-02 current capacity]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-T01 | Power trace width on VLED rail: measure from top layer photograph (calibrated image scale) | Optical / caliper | ≥ 2.0 mm on all VLED/GND power bus segments |  |  |
| FAI-T02 | Thermal imaging: run all 10 strings in one zone at full rated current (180 mA/string) for 60 s; capture IR image | FLIR or equivalent IR camera | No hotspot > 62°C junction temperature; uniform temperature gradient across strings |  |  |
| FAI-T03 | Trace continuity on all 20 power pins: measure resistance pin-to-string LED | 4-wire measurement | Per-trace resistance < 50 mΩ for any signal trace, < 10 mΩ for power rail |  |  |

3b. LED Driver Current Verification  [RISK-05 current hogging · RISK-07 BCR421W]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-E01 | [BLOCKING] Per-string current: measure actual current in each of 10 strings for one wavelength at rated RSET | Clamp meter / shunt, each string individually | All 10 strings within ±5% of target (e.g. 180 mA ± 9 mA). Any string outside ±10% → investigate driver |  |  |
| FAI-E02 | Across-wavelength match: compare 660 nm string mean vs 808 nm string mean current | As above, both wavelengths | Means within ±3% of each other at same RSET (same driver part) |  |  |
| FAI-E03 | BCR421W thermal rise: measure IC case temperature at rated current for 60 s (no heat sink) | NTC or contact thermocouple | Case temperature ≤ 85°C. Above 80°C → evaluate thermal derating |  |  |

3c. Reference Photodiode Optical Isolation  [RISK-06 photodiode contamination]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-P01 | [BLOCKING] Dark-field isolation test: cover all LEDs with opaque black tape; run photodiode readout in synchronous lock-in mode | Bench test, firmware lock-in mode | Photodiode ADC output < 1% of full-scale in absence of LED illumination |  |  |
| FAI-P02 | Response test: uncover LEDs; run all 10 strings in zone at 25% duty cycle; verify photodiode responds | Bench test | Photodiode output increases ≥ 90× compared to dark reading; signal stable across 10 s |  |  |
| FAI-P03 | LED aging simulation: attenuate LED output 30% with ND filter over PDMS window; verify photodiode detects reduction | ND filter + bench test | Reported J/cm² decreases by 28–32% (within ±2% of filter attenuation) |  |  |
| FAI-P04 | PDMS window optical transmission: measure before and after assembly with integrating sphere or photodiode | Integrating sphere / benchtop | ≥ 92% at 660 nm; ≥ 94% at 808–830 nm |  |  |

3d. PDMS Window Adhesion  [RISK-04 CTE delamination]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-M01 | Visual inspection of PDMS–PI bond line: no delamination, no air bubbles > 0.5 mm, no discolouration | Visual, 10× magnifier | No defects visible. Any delamination ≥ 1 mm → reject |  |  |
| FAI-M02 | Peel test (witness coupon from same bonding batch as production units): 90° peel test on 25 mm × 25 mm PDMS–PI bonded coupon. Coupon must be processed through full SiO₂ sputter + O₂ plasma + cure sequence identically to production FPCs. | Instron or similar, 90° peel | Peel force ≥ 150 N/m (per NP-HW-FPC-001 §9.3 and IPC-TM-650 2.4.9). Target ≥ 400 N/m. 100–149 N/m → conditional accept, review bonding process, re-run 3 coupons. < 100 N/m → reject batch; halt production; SiO₂ interlayer process non-conformant — escalate to Manufacturing Eng. |  |  |
| FAI-M03 | Accelerated thermal cycling screen (production batch check — 5 cycles): 5 cycles −20°C to +70°C on 1 witness coupon per production batch. This is a fast screen only — it does NOT replace the full 200-cycle IEC 60068-2-14 process qualification (see §3e FAI-TC items, which must pass before any production begins). | Thermal chamber | No new delamination (visual + 10× magnifier) after 5 cycles. If any delamination: halt batch; run full peel test per FAI-M02. If peel force drops below 150 N/m: reject batch, re-qualify process. Note: FAI-M03 pass alone does NOT confirm long-term reliability — §3e qualification must already be on file. |  |  |

Test rationale: no published fatigue data exists for PI–PDMS bonds under LED-junction thermal cycling conditions (PDMS CTE ≈ 300 ppm/°C vs. PI ≈ 18 ppm/°C, 15:1 mismatch). 200-cycle IEC 60068-2-14 qualification generates this data. If the SiO₂ interlayer process requires adjustment, it must be identified here — not after tooling is committed.

This section covers the one-time process qualification test required before any production zone modules are built. It must be completed and signed off before zone module tooling is released (NP-COORD-001 gate G1-08). The test validates that the SiO₂-interlayer bonding process (NP-HW-FPC-001 §9.2) can survive the 1,800+ thermal cycles expected over device lifetime. FAI-TC02 is BLOCKING — production cannot start without a pass record on file.

3e.  PDMS Thermal Cycling Process Qualification  [RISK-04 — one-time, pre-production]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-TC01 | Bonding process documentation check: confirm SiO₂ sputter parameters (75 nm, RF magnetron), O₂ plasma parameters (100 W / 30 sccm / 60 s), cure (120°C / 2 hr / 50 g/cm²), and < 10 min delay between plasma and bond contact are documented in supplier process control plan before any qualification samples are made. | Process control plan review | All 5 parameters documented and confirmed. Any deviation from NP-HW-FPC-001 §9.2 must be approved in writing by HW EE before qualification proceeds. |  |  |
| FAI-TC02 | [BLOCKING] Full thermal cycling qualification: IEC 60068-2-14 condition Na. 10 bonded PDMS–FPC coupons (same dimensions and process as production). 200 cycles, −10°C to +65°C, 15 min dwell at each extreme, ramp rate ≤ 3°C/min. Inspect 5 coupons at 100 cycles (interim check); inspect all 10 at 200 cycles. | IEC 60068-2-14 thermal chamber; 200 cycles; interim check at 100 cycles | Zero delamination (visual + 10× magnifier) at 100-cycle interim check. Zero delamination at 200 cycles on all 10 coupons. Any delamination at any point → HALT: SiO₂ interlayer process non-conformant. Review sputter thickness, plasma parameters, and cure conditions before re-run. Production CANNOT start until this test passes. |  |  |
| FAI-TC03 | Post-cycling peel test: 90° peel test on all 10 coupons immediately after FAI-TC02 thermal cycling completion. Same test method as FAI-M02 (IPC-TM-650 2.4.9). | 90° peel test, Instron or similar | Peel force ≥ 150 N/m on all 10 coupons after 200 cycles. Target ≥ 400 N/m (pre-cycling target applies here too — cycling should not degrade peel force below 150 N/m). Record min/max/mean — document in process qualification report. |  |  |
| FAI-TC04 | Post-cycling optical transmission: measure optical transmittance at 660 nm and 808 nm through each of the 10 coupons after thermal cycling, using the same spectrophotometer as pre-cycling baseline measurement. | Spectrophotometer, 660 nm and 808 nm | Transmittance ≥ 90% of pre-cycling baseline value at both wavelengths. Any coupon below 90% → investigate for delamination-induced air gap or PDMS haze formation. Reject if < 85% absolute transmittance. |  |  |
| FAI-TC05 | LED junction thermal cycling simulation (supplemental, recommended): assemble 3 complete zone modules with qualified PDMS process; run 500 LED-on/LED-off thermal cycles at maximum rated current (PWM 25% duty, 180 mA peak). Monitor PD1 output each cycle — normalised to initial value. | Production zone module + bench power supply; 500 LED cycles | PD1 output remains within 5% of initial value after 500 cycles (no air gap forming under LED thermal cycling). This test is recommended, not mandatory for production release, but must be completed before claiming > 2-year PDMS reliability. |  |  |
| FAI-TC06 | Qualification report sign-off: all FAI-TC01 through FAI-TC04 results recorded in PDMS Process Qualification Report (document number NP-QR-PDMS-001). Report must be signed by Manufacturing Eng and HW EE before production start. | Document review | Signed qualification report on file. Report includes: supplier name, process control plan ref, test dates, all peel force and transmittance data, thermal chamber calibration certificate. Report retained for device lifetime + 2 years (regulatory retention requirement). |  |  |

SECTION 4 — SHELL INSTALLATION AND ALIGNMENT VERIFICATION

RISK-12 primary mitigation. All 5 zone modules must be verified individually. The ±0.5 mm system alignment requirement is achieved via ±0.3 mm shell tooling tolerance (±0.2 mm margin).

4a. Zone Module Positional Accuracy  [RISK-12 alignment tolerance]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-A01 | [BLOCKING, PENDING REG-1] T1-B (EEG-carrying) module CMM measurement at the socket assigned to 10-20 site Fp1/Fp2 (single shared socket — Fp row holds one tile, REG-1 open item) | CMM / optical profilometer | LED/electrode centroid within ±0.3 mm of the socket's nominal 10-20 target position — socket assignment per docs/np_hex_zm_001.md §3.2/§4a, PROVISIONAL pending gate REG-1 |  |  |
| FAI-A02 | [BLOCKING, PENDING REG-1] T1-B (EEG-carrying) module CMM measurement at the socket assigned to 10-20 site F3/F4 | CMM / optical profilometer | LED/electrode centroid within ±0.3 mm of the socket's nominal 10-20 target position — socket assignment per docs/np_hex_zm_001.md §3.2/§4a, PROVISIONAL pending gate REG-1 |  |  |
| FAI-A03 | [BLOCKING, PENDING REG-1] T1-B (EEG-carrying) module CMM measurement at the socket assigned to 10-20 site C3/C4 | CMM / optical profilometer | LED/electrode centroid within ±0.3 mm of the socket's nominal 10-20 target position — socket assignment per docs/np_hex_zm_001.md §3.2/§4a, PROVISIONAL pending gate REG-1 |  |  |
| FAI-A04 | [BLOCKING, PENDING REG-1] T1-B (EEG-carrying) module CMM measurement at the socket assigned to 10-20 site P3/P4 | CMM / optical profilometer | LED/electrode centroid within ±0.3 mm of the socket's nominal 10-20 target position — socket assignment per docs/np_hex_zm_001.md §3.2/§4a, PROVISIONAL pending gate REG-1 |  |  |
| FAI-A05 | [BLOCKING, PENDING REG-1] T1-B module CMM measurement at the socket assigned to Oz (required for the photoparoxysmal visual-stim safety interlock, docs/np_hex_zm_001.md §4a) | CMM / optical profilometer | Electrode centroid within ±0.3 mm of the Oz socket's nominal target position — PROVISIONAL pending gate REG-1 |  |  |
| FAI-A06 | Inter-electrode spacing: confirm T1-B socket positions match 10-20 system inter-electrode distances | CMM / derived from A01-A05 measurements | All inter-electrode distances within ±0.5 mm of 10-20 system nominal — PROVISIONAL pending gate REG-1 |  |  |
| FAI-A07 | Hex-tile module stiffener datum features (reference holes or edges): verify present and within ±0.1 mm of nominal | CMM | Datum features present; position ±0.1 mm |  |  |
| FAI-A08 | Hex-tile module re-insertion repeatability: remove and re-insert each hex-tile module 3×; re-measure CMM position | CMM, 3 repeats per zone | Position repeatability ±0.1 mm (module self-locates to same position each insertion) |  |  |
| FAI-A09 | Universal orientation key verification: attempt to insert a hex-tile module rotated 60°/120°/180°/240°/300° from correct orientation into a socket | Physical manual test — 5 wrong-orientation attempts per sampled socket | Zero incorrect-orientation insertions possible at any of the 5 rotations. Key geometry alone (no vision, no reading, no firmware) must stop the attempt before electrical contact — verifies the blind-safe claim in docs/np_hex_zm_001.md §4a. |  |  |
| FAI-A10 | Correct-orientation insertion: confirm a module in the correct orientation seats fully and makes electrical contact on the first attempt | Physical manual test, sampled sockets | Full seating and contact confirmed on first attempt at every sampled socket. No false resistance at correct orientation. |  |  |
| FAI-A11 | UID auto-inventory verification: insert a known module type at a sampled socket; confirm np_module_map reports the correct type via UID-based auto-inventory (no ZONE_ID resistor — that mechanism is retired) | Hub MCU diagnostic mode / app inventory view | Reported module type at the socket matches the physical module inserted, within the poll interval specified in docs/np_hex_zm_001.md §4 |  |  |
| FAI-A12 | [BLOCKING] Type-mismatch rejection: attempt to run a protocol whose module map excludes the module type actually inserted at a required socket (e.g. T1-A where T1-B is required); confirm the app blocks the protocol before session start | Live app + firmware test — np_module_map_check_placement() | Session blocked before any stimulation/PBM enable. Unmet socket + required type reported to the user via the app. |  |  |
| FAI-A13 | [BLOCKING] Placement-check software gate: configure a protocol whose module map requires an EEG-capable module at a given socket; leave that socket unpopulated or populate it with a non-EEG module; confirm the app blocks the protocol and identifies the unmet socket | Live app + firmware test — np_module_map_check_placement() | Protocol is disabled with the specific unmet socket identified to the user. Protocol becomes runnable immediately once a matching module is inserted at that socket — no other action required. |  |  |
| FAI-A14 | Mixed-type insertion: populate a representative cluster with a mix of T1-A/B/C module types in arbitrary sockets; confirm all types insert, seat, and are correctly identified regardless of socket position | Physical + firmware test | Every type-agnostic combination inserts and is correctly identified. No socket position is type-restricted, per docs/np_hex_zm_001.md §4a "any type inserts into any socket." |  |  |
| FAI-A15 | Eject lever accessibility test (RISK-22): 5 test subjects with diagnosed Parkinson's disease (Hoehn & Yahr scale II–III) or post-stroke hand weakness (grip strength ≤ 8 kg). Each subject removes and reinserts all 5 zone modules twice (10 total operations). No coaching; standard user instructions only. | HFE formative test — 5 subjects with Parkinson's/post-stroke (pre-production) | All 5 subjects complete each extraction in ≤ 10 s with no more than 2 attempts. No reported pain or excessive effort. Lever detent must not prevent extraction. Lever must not accidentally eject during normal wear simulation (5 min handling without deliberate ejection attempt). |  |  |
| FAI-H01 | HRV coherence baseline (standalone): subject performs 5-minute guided breathing session at 6 breaths/min. App displays real-time coherence score. Verify: PPG R-R interval successfully detected from auricular clip (>95% beat detection); coherence score updates at ≤2s interval; bone conduction audio breathing cue audible at ≥60dB at ear. | App + firmware verification | Beat detection ≥95% (validated against ECG reference in bench test). Coherence score range 0–10 correctly calculated. Audio cue timing: ±0.1s of target. |  |  |
| FAI-H02 | HRV+taVNS synchronization: 10-minute session, HRV+taVNS synchronized protocol. Verify firmware correctly identifies inspiration phase from PPG waveform and triggers taVNS pulse within ±150ms of peak inspiration for ≥80% of respiratory cycles. Subject: 3 healthy adults. | 3 subjects, bench + human verification | Inspiration-triggered stimulation latency ≤150ms in ≥80% of cycles. No taVNS stimulation during expiration-only phases. No subject reports unexpected sensation from mis-timed stimulation. |  |  |
| FAI-H03 | HRV+EEG dual biofeedback: 10-minute session, both coherence score and EEG alpha/theta band power displayed simultaneously. Verify: both metrics update independently without interference; EEG signal quality unchanged when taVNS is running concurrently (signal-to-noise check). | App + signal quality verification | Both metrics update at ≤2s intervals. EEG SNR during concurrent taVNS ≥90% of baseline (taVNS artefact <5µV RMS after filtering). |  |  |
| FAI-H04 | HRV+PBM multi-modal session: 15-minute combined session (PBM zone 1–5 at standard dose + HRV coherence training simultaneously). Verify: PBM dosimetry unaffected by concurrent HRV session; HRV coherence scores in expected range (≥5.0 for subjects trained in resonance breathing); no thermal interaction between PBM NTC monitoring and PPG signal processing. | 3 subjects with prior HRV biofeedback experience | PBM J/cm² dose within ±5% of target. HRV coherence ≥5.0 in all 3 subjects (indicates usable biofeedback loop). NTC temperature reporting unaffected. |  |  |
| FAI-HD01 | sLORETA source localization accuracy: T2 qEEG session (21-ch, eyes-closed resting state, 5 min). sLORETA algorithm computes cortical source map. Verify: source reconstruction completes within 10s; DLPFC peak activation localises to within ±15mm MNI coordinates of expected region (validated against published normative qEEG atlas). Bench test with 5 healthy adults. | Computational + bench verification (5 subjects) | sLORETA output within 10s. DLPFC localization error ≤15mm MNI vs atlas. Consistent across 5 subjects (σ ≤ 8mm inter-subject). |  |  |
| FAI-HD02 | HD-tDCS electrode mapping: verify firmware correctly translates sLORETA MNI target coordinates to 10-20 electrode positions and generates 4×1 montage configuration. Test with 5 simulated target locations (F3/Fz/Cz/P3/Oz). Confirm current distribution computed correctly for each montage. | Firmware + computational verification | Correct 10-20 position identified for all 5 test targets. 4×1 montage configuration matches manual calculation ±5%. Current sum across all electrodes = 0 (charge-balanced) in all montages. |  |  |
| FAI-HD03 | HD-tDCS focality vs standard tDCS: bench measurement using gel phantom. Compare electric field distribution for (a) standard 2-electrode 2mA tDCS and (b) 4×1 ring HD-tDCS 2mA at same scalp location. Measure with embedded electrode array (or finite element model validated against phantom). Confirm HD-tDCS half-maximum field width ≤ 3× narrower than standard. | Phantom / FEM bench test | HD-tDCS FWHM ≤ 3cm vs standard tDCS FWHM ≥ 8cm in gel phantom (consistent with Bikson lab published data). Peak field strength within safety envelope. |  |  |
| FAI-HD04 | Simultaneous EEG + HD-tDCS signal quality: 21-ch EEG recorded during and after 1mA HD-tDCS (3 min on, 2 min post). Verify: (a) EEG artefact during stimulation is rejected correctly by artefact filter; (b) post-stimulation EEG SNR recovers to ≥95% of pre-stimulation baseline within 30s (electrode polarisation recovery); (c) no electrode damage or impedance drift >5kΩ after 10 consecutive sessions. 3 bench subjects. | 3 subjects, signal quality + electrode durability | EEG recovers ≥95% SNR within 30s post-stimulation. Impedance drift ≤5kΩ after 10 sessions. No visible electrode damage on inspection. |  |  |
| FAI-CV01 | Cervical VNS electrode placement verification: confirm correct gel electrode placement on skin overlying sternocleidomastoid muscle / carotid sheath using anatomical landmarks. 5 subjects (3 male, 2 female, BMI range 20–32). Verify impedance <5kΩ at placement; no stimulation over carotid artery directly (ultrasound-guided confirmation in bench study). | Anatomical bench verification, 5 subjects | Impedance <5kΩ at correct placement in all 5 subjects. Ultrasound confirms electrode not directly over carotid bifurcation. Placement guide in IFU confirmed adequate by all 5 subjects. |  |  |
| FAI-CV02 | Cardiac safety interlock: connect cervical VNS to simulated artefact ECG waveform generator. Deliver 1mA cervical stimulation. Inject simulated HR change of +16 BPM at t=3s. Verify: safety MCU detects HR change >15 BPM within ≤500ms and cuts stimulation enable line. No stimulation resumes without user confirmation. | Bench safety interlock test (simulated ECG generator) | Stimulation cutoff within ≤500ms of HR change >15 BPM threshold. Stimulation does not resume automatically. Log entry created in SHDR with timestamp and HR value at cutoff. |  |  |
| FAI-CV03 | Cervical VNS tolerability: 10-minute session at 0.5mA (low) and 1.0mA (standard), bilateral cervical placement. 5 healthy adult subjects. Collect: subject comfort rating (NRS 0–10), carotid palpation result before/after, ECG rhythm pre/during/post. No vasovagal response, no carotid bruit change, no arrhythmia. | 5 subjects, tolerability + cardiac monitoring | NRS comfort ≤4/10 at 1.0mA in all subjects. No vasovagal episode. No cardiac arrhythmia on ECG during or post-stimulation. No carotid bruit change on auscultation. |  |  |

4b. FPC Routing Path Verification (in-shell)  [RISK-11 bend radius]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-R01 | Zone 1 FPC: visual inspection of routing path in assembled headset — no visible kink, tight bend, or pinch point | Visual + endoscope (if needed) | No kink visible; FPC follows channel smoothly |  |  |
| FAI-R02 | Zone 2 FPC: same as FAI-R01 (highest risk — compound frontal curvature) | Visual + endoscope | No kink; smooth channel follow |  |  |
| FAI-R03 | Zones 3–5 FPC: visual inspection | Visual | No kink on any zone |  |  |
| FAI-R04 | NEEDS REVIEW — was scoped to the retired 5-zone-slot FPC routing architecture (RISK-11/RISK-17). The hex-tile/socket architecture (NP-HEX-ZM-001) per-socket interconnect this test would run against is not yet designed (see MECH-1/MECH-2/SMART-1 in docs/np_hex_zm_001.md §7). Cannot be executed as written. | Radius gauge, bend check tool | NEEDS RE-SCOPE once hex-tile interconnect design exists |  |  |
| FAI-R05 | Zone 1 insertion/removal cycle test: insert/remove 10×; inspect FPC for crease, whitening, or deformation | Manual test + visual | No crease, whitening, or deformation after 10 cycles |  |  |
| FAI-R06 | Cross-zone interference: remove Zone 1 module; confirm Zone 2 FPC is not disturbed | Visual | Zone 2 FPC remains in channel during Zone 1 removal |  |  |
| FAI-R07 | ZIF lever access: confirm user can actuate ZIF lever (full 180° flip) without contacting FPC with fingertip | Manual test, 3 operators | All 3 operators actuate lever without FPC contact |  |  |

Self-sealing co-molded silicone gasket — no user action, no RTV, no re-sealing procedure. FAI-IPX-02 is BLOCKING.

4c.  Sealing and IPX4 Verification  [RISK-16 — Option A gasket]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-IPX-01 | Initial IPX4 baseline: assembled headset, all 5 zone modules factory-installed. IEC 60529 IPX4 water spray from all directions, 10 minutes. | IEC 60529 IPX4 nozzle, 10 min | No water ingress to Hub PCB, any connector, or zone module PCB. EEG impedance all channels < 10 kΩ after test (functional check). |  |  |
| FAI-IPX-02 | [BLOCKING] Field-replacement IPX4 re-test: remove and re-insert all 5 zone modules 10 times each (50 total swaps, no tools, no RTV, user-pace); then repeat full IEC 60529 IPX4 spray test without any re-sealing. | 10 swap cycles per zone then IEC 60529 IPX4 | Same pass criterion as FAI-IPX-01. Zero water ingress after 10 field swaps. Any ingress → HALT production; gasket geometry or shell rim tolerance non-conformant. Escalate to Mechanical Engineering immediately. |  |  |
| FAI-IPX-03 | Gasket visual and dimensional check after 10 swap cycles: inspect all 5 zone module gaskets under 10× magnification; measure gasket cross-section height on 2 modules with calibrated comparator. | 10× magnifier; calibrated comparator (2 modules) | No tears, cuts, or embedded debris. Compression set < 10%: gasket height ≥ 1.80 mm (vs. 2.00 mm nominal) after cycle release. |  |  |
| FAI-IPX-04 | Gasket shelf-life / ageing qualification (first-article qualification, not every lot): store 5 sealed zone modules at 40°C / 75% RH for 90 days (equivalent to ~2 years at ambient per Arrhenius). Perform IPX4 test and gasket dimension check at end of storage. | IEC 60068-2-66 damp heat storage; then IPX4 + dimensional check | IPX4 pass. Gasket dimensions within 5% of nominal. Compression set < 15% after ageing. Qualification pass required before first production run. |  |  |

SECTION 5 — CONNECTOR LIFECYCLE VALIDATION

Zone modules are user field-replaceable. The Hirose FH34S connector is rated ≥1,000 cycles. FAI validates first-article connector performance at incoming (10-cycle) and production-release (50-cycle) levels.

5a. Connector Insertion Cycle Test  [RISK-01 connector family · RISK-11 routing path]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-K10 | Baseline contact resistance (before cycling): measure all 20 contacts | 4-wire, all 20 pins | All pins < 50 mΩ |  |  |
| FAI-K11 | Insert/remove zone module 50× (simulated field replacement pace — one cycle per 15 s) | Manual cycling, count logged | 50 cycles completed without mechanical failure of lever, body, or contacts |  |  |
| FAI-K12 | Post-cycle contact resistance: measure all 20 contacts | 4-wire, all 20 pins | All pins < 50 mΩ. Increase > 20 mΩ from baseline → flag for investigation |  |  |
| FAI-K13 | Post-cycle visual: inspect connector body and FPC tail for wear marks, gold wear-through, or fretting | Visual, 10× magnifier | No exposed base metal on contact pads; no cracking of FPC coverlay at tail |  |  |
| FAI-K14 | Post-cycle FPC position accuracy (re-run FAI-A01 to FAI-A05 on Zone 1 after 50 cycles) | CMM | Zone 1 position still within ±0.3 mm after 50 insertion cycles |  |  |

SECTION 6 — FUNCTIONAL SYSTEM TEST (ASSEMBLED HEADSET)

Run after all physical inspection items pass. Requires firmware-loaded headset connected to app.

6a. End-to-End Zone Function  [RISK-05 · RISK-06 · RISK-12]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-SY01 | All populated sockets illuminate at 25% duty cycle when the "All" zone (protocols/predefined/00-zones.npps) is activated — app confirms socket-level active status, photodiode returns nonzero dose reading per populated socket | App UI, firmware telemetry | Every populated socket reports active; J/cm² > 0 for each |  |  |
| FAI-SY02 | Zone-selective deactivation: activate the "Frontal Left" zone only (protocols/predefined/00-zones.npps); confirm its member sockets are active and all sockets outside it remain inactive | App UI | Correct zone-membership isolation confirmed against the 00-zones.npps socket list |  |  |
| FAI-SY03 | PDMS fouling simulation: apply light smear of petroleum jelly to Zone 1 PDMS window; confirm photodiode detects reduced output and app flags maintenance reminder | App UI + petroleum jelly test | App shows reduced dose for Zone 1; maintenance reminder displayed within 2 sessions |  |  |
| FAI-SY04 | EEG signal quality on all 8 electrodes: confirm impedance < 10 kΩ with hydrogel tips installed | App impedance display | All 8 electrodes < 10 kΩ (green) |  |  |
| FAI-SY05 | Closed-loop session: run 5-minute EEG-adaptive PBM session; confirm adaptation events logged in UHDR | App session log | At least 1 adaptation event logged; UHDR session record created and encrypted |  |  |

SECTION 7 — NON-CONFORMANCE SUMMARY

List all items that did not pass during this FAI. All non-conformances must be dispositioned before production release — either by corrective action (re-inspect after fix) or by formal Engineering deviation approval. No BLOCKING item may be waived.

| FAI Item | Non-Conformance Description | Disposition | Corrective Action Ref | Closed by / Date |
| --- | --- | --- | --- | --- |

SECTION 8 — FAI RESULT AND SIGN-OFF

All checklist items in Sections 1–6 must show PASS, or non-conformances in Section 7 must be fully closed, before production release is authorised.

| Overall FAI result: | PASS / FAIL | Build lot approved for production: | YES / NO |
| --- | --- | --- | --- |
| Manufacturing Engineer: |  | Date: |  |
| Hardware EE Lead: |  | Date: |  |
| Quality / Reliability: |  | Date: |  |
| Program Manager: |  | Date: |  |
| Notes: |  |  |  |

SECTION 9 — RISK MITIGATION CROSS-REFERENCE

All FAI items trace back to at least one risk register entry or specification requirement.

| Risk ID | Risk Title (abbreviated) | FAI Items | Status |
| --- | --- | --- | --- |
| RISK-01 | Connector family — Hirose FH34S | FAI-D05, FAI-K01–FAI-K04, FAI-K10–FAI-K14 | MITIGATED |
| RISK-02 | FPC power trace current capacity | FAI-T01, FAI-T02, FAI-T03 | MITIGATED |
| RISK-04 | PDMS–PI CTE mismatch delamination | FAI-M01, FAI-M02, FAI-M03, FAI-P04 | OPEN (process qualification pending) |
| RISK-05 | LED current hogging (per-string CCC) | FAI-E01, FAI-E02, FAI-E03, FAI-SY01 | MITIGATED |
| RISK-06 | Reference photodiode optical isolation | FAI-P01, FAI-P02, FAI-P03, FAI-SY03 | MITIGATED |
| RISK-07 | BCR421U vs BCR421W variant | FAI-C01, FAI-C02, FAI-D04 | MITIGATED |
| RISK-08 | LED Vf binning not on PO | FAI-L01–FAI-L05, FAI-D08 | MITIGATED |
| RISK-10 | ED copper default on FPC fab | FAI-D06, FAI-D07, FAI-F01, FAI-F02 | MITIGATED |
| RISK-11 | Shell bend radius not specified | FAI-R01–FAI-R07, FAI-K14 | MITIGATED |
| RISK-12 | Zone alignment tolerance not validated | FAI-A01–FAI-A08, FAI-SY04 | MITIGATED (this FAI) |
| RISK-15 | Zone keying — eliminated by hex-tile/socket architecture (NP-HEX-ZM-001); universal orientation key + software placement-check | FAI-A09, FAI-A10, FAI-A11, FAI-A12, FAI-A13, FAI-A14 (all corrected) | MITIGATED (this FAI, corrected) |
| RISK-16 | IPX4 field-replacement — Option A gasket | FAI-IPX-01, FAI-IPX-02, FAI-IPX-03, FAI-IPX-04 | MITIGATED (this FAI) |
| RISK-22 | Zone module extraction — reduced grip / tremor (sliding eject lever) | FAI-A15 | MITIGATED (this FAI) |
| RISK-04 | PDMS–PI CTE mismatch thermal cycling delamination | FAI-M01, FAI-M02, FAI-M03, FAI-TC01–FAI-TC06 | MITIGATED (this FAI) |

SECTION 10 — REVISION HISTORY

| Rev | Date | Author | Description |
| --- | --- | --- | --- |
| A | 2026-05-06 | NeurOne Engineering | Initial release. RISK-12 primary mitigation (alignment FAI). Consolidates FAI items from RISK-01, 02, 04, 05, 06, 07, 08, 10, 11, 12. 6 sections, 50+ inspection items. |