NeuroPulse

Master Design Brief — Revision 3

Closed-Loop Multi-Modal Neuromodulation Platform

May 4, 2026  ·  Revision 3 — incorporates all decisions through current session

# Revision 3 — Key Additions Since Revision 2

| Category | New decisions added |
| --- | --- |
| UHDR / SHDR data architecture | Complete UHDR/SHDR terminology framework. SHDR linked to device ID + warranty owner only — never user identity. UHDR encrypted with user-held key. Separate eMMC partitions from first firmware line. All boundary cases resolved by derivation rule. Session count in SHDR is unsigned integer with no user link. |
| Clinical consent engine | Use case library (9 defined use cases). Minimum necessary data mapper. Plain-language consent document generator (not legal forms). Subscription tiers (Monitor $49, Assess $149, Full Clinical $299, Research $599/mo). Expansion workflow with differential consent and 'approve with questions' path. |
| Research consent + partnership | Three a priori consent layers (contact, category, blanket). 9 research categories. Blanket consent still receives per-study engagement notifications. Results opt-in (plain-language summary + paper link + null result notifications). Study suggestion portal (votes, participation intent, pledges). Crowdfunding mechanics (pledge → pilot → grant catalyst). POA workflow with human document review. Researcher portal with pre-identified cohort data. |
| Durability engineering | ITO → AgNW lens coating. Hard clamshell case replaces pouch. Intranasal probe hub dock. Palladium fabric (replaces silver — marketing differentiator + fleet-verified claim). N52 → N42 magnets. 1,000-cycle zone connectors. Interface protection covers (all tethered). Braided aramid USB-C cable. MagSafe hard gold contacts. 22F supercapacitor. IPX4 target. Audio cup user-replaceable mesh frame + mesh cleaning brush. |
| Calibration + self-maintenance | Reference photodiode per zone (eliminates 3-year service calibration visit). Plasma-activated anti-fouling PDMS. Fluxgate zero-field nulling + geomagnetic comparison. Hub NTC cross-calibration of headset NTCs. ADS1299 internal reference self-calibration at session start. EC lens transition time monitoring for contact resistance. Audio cup driver impedance monitoring for mesh fouling. Probe tip reference LED at probe base. |
| Software resilience | Dual-bank OTA firmware (Bank A runs while Bank B staged, cryptographic verification before switch). USB-C DFU recovery mode. Safety MCU firmware only updated in maintenance mode. Industrial eMMC + LittleFS. Firmware partition write-protected. Automated nightly incremental UHDR backup. Cloud service commitment in warranty. 7-day OS compatibility SLA. New device migration prompt. |
| Service network | Tier A optical centers (S3 Rx + lens + calibration, 4-hr cert, $400–600 equipment). Tier B electronics repair (FPC, DFU, data recovery, 6-hr cert). Tier C retail triage (Best Buy Geek Squad type, 2-hr cert, no equipment). NeuroPulse depot (backstop + T2 loaner). Sliding rail lens mount (user self-install). Lever ZIF connectors (user zone module swap). Tool-free hub fan. Design changes reduce service dependency substantially. |
| Box contents (updated) | Hard clamshell case (replaces pouch). Braided aramid USB-C cable (spare included). 45W GaN charger (branded NeuroPulse). Mesh cleaning brush. Boa replacement cable + hook tool. Moisture-barrier electrode tip hydration caps. Humidity indicator card. Pre-impregnated cleaning cloth packets. Interface covers (installed + one spare set each type). |

# 1. Executive Summary

| NeuroPulse Home — Tier 1 Consumer wellness · FDA-exempt · 8 modalities · $449–$1,199 - 600 LEDs (300×660nm + 300×810nm) on FPC zone strips — uniform field, 80,000-hr lifetime - 400 mW/cm² peak pulsed irradiance (firmware-only, $0 BOM) — matches Vielight Neuro Pro 2 headline spec - 8-channel semi-dry EEG, autonomous closed-loop EEG-adaptive stimulation without phone - 5-layer EMF shielding + active Helmholtz cancellation — only consumer brain wearable with measured shielding - Pricing: $449 Core · $599 Lite · $849 Standard ★ · $1,199 Premium - No mandatory subscription. All core functions offline-capable permanently. | NeuroPulse Pro — Tier 2 Clinical/research · FDA 510(k) target · all modalities · $4,999–$13,999 - 21-channel qEEG wet gel: full 10-20 + FC3/FC4 (M1 TMS targeting) + Oz (photoparoxysmal) + A1/A2 (normative database reference) - TMS focal figure-8 coil (rTMS/TBS) · 1170nm deep PBM (35–40mm subcortical) · clinical tACS ≤4mA - HIPAA cloud · EHR (FHIR R4) · multi-patient dashboard · sLORETA source imaging · LSL streaming - Pricing: $4,999 Pro Entry · $13,999 Pro Full · $1,800/yr service (same-day loaner) - CPT 90867–90869 TMS reimbursement · 510(k) pathway 18–36 months · $2–5M budget |
| --- | --- |

# 2. Data Architecture — UHDR / SHDR Framework

| Record type | Definition, contents, access, and storage |
| --- | --- |
| UHDR — User Health Data Record (user's property, never accessed by NeuroPulse) | DEFINITION: Any record describing the user's biological state, cognitive function, therapeutic response, or health outcomes. Defining test: does this record tell us something about the person? If yes → UHDR. CONTENTS: EEG waveforms (all channels) · HRV time series · PPG optical signal · Neurofeedback performance scores · Session timestamps and duration · Protocol parameters used per session · Closed-loop adaptation events · PBM dose (J/cm²) per zone · User-entered symptom/outcome logs · Eye-open/closed state during sessions. ACCESS: User (always, unconditionally). Clinicians (per-element, per-use-case consent, time-limited, audited, revocable). Researchers (anonymised aggregate, separate IRB + explicit research consent). NeuroPulse: NEVER. STORAGE: On-device eMMC UHDR partition, AES-256 encrypted with user biometric-derived key. NeuroPulse does not hold decryption key. Optional E2E encrypted cloud backup (user-held key). Automatic incremental backup on USB-C connect. |
| SHDR — System Health Data Record (NeuroPulse property, never linked to user identity) | DEFINITION: Any record describing device performance, wear state, and operational history — independent of who is wearing it. Defining test: does this tell us something about the device's condition, containing nothing that reveals user biology? If yes → SHDR. CONTENTS: LED output ratio per zone · NTC temperature profiles · EMF shielding attenuation ratio · Device session count (unsigned integer, no timestamps) · Consumable session counts per component · USB-C insertion counter · PD negotiation log · Impact events (g-force, orientation — between sessions only) · Fan RPM · Supercapacitor cycles · Firmware version history · OTA success/failure log · Accessory authentication pass/fail · Calibration coefficient history. ACCESS: NeuroPulse (fleet database, maintenance prediction, reliability improvement). Warranty owner (maintenance alerts derived from SHDR). Researchers (aggregate only, per data sharing agreement). LINKAGE: SHDR is linked to device ID + warranty owner ID only. Never to user identity. The SHDR database and UHDR storage share no common identifier — they cannot be joined. STORAGE: On-device eMMC SHDR partition, separate encryption from UHDR. Uploaded to NeuroPulse fleet database on USB-C connect (warranty consent required). |
| Boundary case resolution rule | When in doubt → classify as UHDR. Reclassification to SHDR requires positive demonstration of no user biology content. SPECIFIC RESOLUTIONS: Raw EEG impedance (reflects user scalp biology) → UHDR. Derived trend slope across sessions (reflects electrode tip condition only) → SHDR. Raw accelerometer during active sessions → UHDR. Impact events detected between sessions → SHDR. Raw ambient light values (reveal when/where user uses device) → UHDR. Cumulative UV exposure index (single dimensionless engineering number) → SHDR. Raw VNS contact impedance → UHDR. Contact resistance trend → SHDR. IR eye state during sessions → UHDR. Safety interlock log (device behaviour record) → SHDR. |

# 3. Clinical Consent Engine

| Subscription tier | Price | Use cases included | UHDR elements accessible (with user consent) | Target clinician |
| --- | --- | --- | --- | --- |
| Monitor | $49/mo/patient | Treatment adherence monitoring · Protocol compliance tracking | Session timestamps · Session duration · Protocol parameters used | Primary care, wellness providers, care coordinators checking adherence between visits |
| Assess | $149/mo/patient | All Monitor + EEG session review · Neurofeedback performance · Treatment efficacy tracking · Sleep session monitoring | Adds: EEG waveforms · Neurofeedback performance scores · PBM dose logs · Eye-state during session | Neurologists, psychiatrists, EEG specialists, sleep clinicians |
| Full Clinical | $299/mo/patient | All Assess + Autonomic monitoring (HRV) · Closed-loop response review · Patient-reported outcomes · TMS response monitoring | Adds: HRV time series · PPG optical signal · Closed-loop adaptation events · User-entered symptom logs (if separately consented) · EEG FC3/FC4 channels specifically for TMS | Tier 2 neuromodulation programmes, TMS clinics, integrative psychiatry, Parkinson's specialists |
| Research | $599/mo/study | IRB-defined custom use case combination (reviewed by NeuroPulse before activation) | IRB-approved minimum necessary set only · Anonymised output · k-anonymity ≥10 required · No individual identifiers | Academic clinical trials · IRB-approved observational studies · Population meta-analysis |

## 3.1 A Priori Research Consent — Four Onboarding Screens

| Layer | Question | If yes | If no | Effect on device |
| --- | --- | --- | --- | --- |
| L1 — Contact consent | Can we contact you about future research opportunities? | Provide preferred contact method + frequency limit. POA holders upload POA document (human review, 3 business days). All research contact goes to POA holder if applicable. | No research contact sent. Category and blanket consents still recorded. Changeable at any time. | None — device fully functional regardless |
| L2 — Category consent | Which research areas would you broadly support? (9 categories: AD/dementia, Depression, PTSD, TBI, Sleep, Attention, Parkinson's, Healthy ageing, Visual health) | Receives per-project contact requests only for studies in selected categories. Each contact is a fresh per-project decision — no data shared until per-project consent given. | Not contacted for that category. Other categories unaffected. Category selections changeable anytime. | None |
| L3 — Blanket consent | Do you pre-approve all NeuroPulse-reviewed research uses of your anonymised data? | Data included in all reviewed studies. STILL receives per-study engagement notifications (not consent requests — the patient is kept informed and can opt out per-study). Anonymisation standard: k≥10, no device/user ID, no sub-weekly timestamps, no sub-regional geography. | Per-category and per-project process applies. Not a lesser choice — equal in NeuroPulse's treatment. | None |
| L4 — Results + community | Would you like to hear study results? Would you like to join the suggestion portal? | Plain-language results notification per contributed study (including null results) + paper link. Access to suggestion/voting/pledge/crowdfunding portal. Suggestions visible to researcher community. | No results contact. No portal access. Both changeable at any time in Settings → Research + Privacy. | None |

# 4. Configurations + Pricing

| Configuration | BOM est. | COGS | Retail | GM% | Included modalities + key items | Target customer + justification |
| --- | --- | --- | --- | --- | --- | --- |
| Core — EEG only | $165 | $255 | $449 | 43% | 4-ch EEG · all connectivity (USB-C/BT/Wi-Fi) · EMF shielding (palladium fabric) · full processor stack · 8GB eMMC · hard clamshell case · interface covers (installed + spare set) · braided aramid cable · 45W GaN charger | Cautious entry buyer. Neurofeedback baseline. B2B modular base unit. Establishes field upgrade path. |
| Home Lite | $260 | $365 | $599 | 39% | All Core + PBM 5 zones (660+810nm, 600 LEDs FPC, 400 mW/cm² peak pulsed) · 8-ch EEG · VNS+HRV auricular clip (A1/A2 EEG references included) | Wellness-focused consumer. First meaningful PBM config. Targets Neuronic/Vielight alternative seekers below $600. |
| Home Standard ★ | $395 | $530 | $849 | 38% | All T1 modalities: PBM 600 LEDs · intranasal bilateral 2λ 3-depth · 8-ch EEG · BES · tDCS · VNS+HRV · audio cups (planar + bone conduction) · standard lens (35% VLT, 108 micro-LEDs/lens) · S1 opaque shade · all covers · hard case | Primary flagship. Biohacker 35–65. All T1 modalities. 50× more transcranial LEDs than Vielight at 17% of price. Primary marketing SKU. Full SHDR fleet telemetry + maintenance predictions. |
| Home Premium | $450 | $610 | $1,199 | 49% | All T1 + EC lens upgrade (5–75% VLT bistable, +$89 value) · 2yr warranty · priority support · NeuroPulse-branded 45W GaN charger | High-income buyer, premium lens day-one. Highest T1 gross margin. EC lens + palladium fabric as premium differentiators. |
| Pro Entry | $820 | $1,350 | $4,999 | 73% | All T1 + 21-ch qEEG wet gel (FC3/FC4/Oz/A1/A2) · 1170nm deep PBM LDs + TEC · clinical tACS ≤4mA 16-ch · HIPAA cloud · EDF+ · sLORETA source imaging | Neurology clinic, concussion rehab, academic PI (IRB exemption seeding). CPT 90867–90869 TMS eligible as add-on. |
| Pro Full | $1,480 | $2,600 | $13,999 | 81% | All T2: Pro Entry + TMS focal figure-8 coil hub · multi-patient dashboard · research mode · LSL streaming · scripting API · FHIR R4 EHR · anonymised session tag for clinical multi-patient · $1,800/yr service | Academic medical centre, psychiatry clinic, neuromodulation research lab. Same-day loaner via service contract. Highest absolute gross margin. |

# 5. Durability + Maintenance Engineering

## 5.1 All design changes — status and impact

| Change | Status | BOM delta | Rationale + tooling timing |
| --- | --- | --- | --- |
| ITO → AgNW (silver nanowire) conductive lens coating | PRE-TOOLING CRITICAL | +$8–12/lens | ITO has 0.5% strain-to-failure — cracks on any flex or point impact. AgNW tolerates 5–10% flex, compatible with hard coat overlay, maintains 85–90% transmission. Must specify before lens tooling is cut. |
| Hard clamshell case replaces soft pouch | PRE-PRODUCTION CRITICAL | +$8–14 | Lens scratching from bag contents certain within first month without case. Moulded case includes headset form, lens support, probe dock, accessories tray. Doubles as shipping container. |
| Intranasal probe hub dock (moulded) | PRE-TOOLING CRITICAL | Hub retool | Probe dropped probe-first fractures Y-junction. Hub dock is only reliable storage. Cannot be retrofitted — must be in hub tooling from day one. |
| Palladium-coated EMF shielding fabric | DECIDED | +$6/headset | Silver tarnishes 12–18 months daily sweat use. Palladium immune to tarnishing for device lifetime. Fleet SHDR shielding attenuation data will demonstrate stable performance over years — measurable, publishable marketing claim. |
| Reference photodiode per zone (behind PDMS window) | PRE-PCB LAYOUT CRITICAL | +$2 total (5 zones) | Detects LED output drift AND window fouling simultaneously. Eliminates 3-year service center calibration visit entirely. Most important single maintenance engineering decision. |
| Braided aramid USB-C cable + dual silicone strain relief | DECIDED | +$3–4/cable | Commodity cables fail at strain relief within 6–18 months. Braided aramid rated 50,000+ flex cycles. Spare in box. Most-handled component in the system. |
| MagSafe hard gold contact faces (>0.5µm cobalt-alloyed) | PRE-PRODUCTION | +$1.20 | Oxidised contacts raise resistance → power throttling during sessions. Hard gold resists wear 10× better than soft gold. 10,000-cycle rating at full current. |
| 22F supercapacitor (from 10F) | DECIDED | +$1.80 | Allows 50% capacitance degradation over 5 years while maintaining transient absorption. Hub NTC monitors temperature for aging estimation. |
| Dual-bank OTA firmware architecture | PRE-FIRMWARE CRITICAL | $0 (software) | Failed update bricking is 100% manufacturer-caused. Bank A runs while Bank B staged. Boot switches only after cryptographic hash verification. USB-C DFU recovery if both banks corrupted. |
| Industrial eMMC (SLC cache, 30,000+ P/E) + LittleFS | COMPONENT SELECTION | +$2.40 | Consumer eMMC (1,000–3,000 P/E) may fail within 3–5 years at daily use. LittleFS manages bad blocks gracefully. Firmware partition write-protected. |
| Zone slot connectors: 1,000-cycle rated (Molex SlimStack) | PRE-PCB LAYOUT CRITICAL | +$2.00 | Standard connectors rated 100–500 cycles — inadequate for modular users. Specify before PCB layout. |
| Interface protection covers (all tethered) | PRE-TOOLING CRITICAL | +$8–9 total | Zone plugs (Shore 30A silicone, 5 colours, IP54). Port covers (TPE + steel disc + silicone fins). Lens rim guards (Shore 85A TPU, clear). Anchor posts moulded into shell at zero cost if specified before first cut. |
| Audio cup user-replaceable snap-in mesh frame | PRE-TOOLING | +$0.60 | Earwax degrades both acoustic and RF shielding functions simultaneously. Snap-in frame enables user cleaning and annual replacement ($9.99/pair). |
| Sliding rail lens mount (user self-install) | PRE-TOOLING | +$1.20 | Eliminates lens alignment jig requirement. User slides damaged lens out, new lens in, clips retention tab. App validates via EC transition + micro-LED test. Eliminates Tier A installation service visits. |
| Lever-actuated ZIF connectors for zone modules | PRE-PCB LAYOUT | Included in connector upgrade | Standard ZIF requires tools and care. Lever-actuated ZIF with supplied lever tool enables user zone module self-service. |
| Automated nightly UHDR backup (incremental) | PRE-LAUNCH SOFTWARE | $0 | When on USB-C power, silently encrypts and pushes new UHDRs to backup destination. Failure becomes hardware swap not data loss. Eliminates most data recovery service events. |
| N52 → N42 magnets in lens rim | DECIDED | −$0.80 | N52 brittle under corner impact. N42 significantly more impact-tolerant with adequate holding force at slightly larger format. |
| ADS1299 internal reference self-calibration at session start | FIRMWARE | $0 | Routes internal reference to all channels, measures gain/offset, applies correction coefficients. Eliminates EEG amplitude drift over device lifetime. Coefficients logged in SHDR. |
| Plasma-activated PDMS (anti-fouling) for optical windows | PRE-PRODUCTION | Minimal | Hydrophilic surface repels sebum fouling. Standard microfluidics treatment. Same material, surface treatment at manufacture. |

# 6. Service Network

| Tier | Examples | Service tasks | Certification | Equipment cost | Revenue/yr at scale | Launch timing |
| --- | --- | --- | --- | --- | --- | --- |
| Tier A — Optical centers | LensCrafters · Pearle Vision · Independent opticians | S3 Rx clip manufacture + fitting (primary) · Lens replacement (standard + EC) · Calibration (reference surface, app-guided) | 4-hr initial · 1-hr annual online | $400–600 loaned calibration reference + $80–120 alignment jig (eliminated by sliding rail — jig becomes optional) | $8K–35K/yr at scale | Year 1 — already engaged via S3 Rx programme. Anchor of entire network. |
| Tier B — Electronics repair | uBreakiFix / Asurion · iFixit partners · Independent repair shops | Zone module FPC swap · Hub fan replacement · DFU firmware recovery · eMMC data recovery · Impact inspection · Fluxgate calibration (Helmholtz coil reference) | 6-hr initial · 2-hr annual practical | Existing ESD workstation + eMMC recovery adapter ($150–200) + NeuroPulse DFU software (downloaded) | $4K–18K/yr at scale | Year 2 — major metros first, expand with fleet. Becomes primary for electronic failures. |
| Tier C — Retail triage | Best Buy Geek Squad · Apple Authorized Service Providers | Warranty intake + triage · Routing to Tier A/B · Consumable sales · DFU if trained | 2-hr initial · 30-min annual online | None — partner app access only. Consumable stock. | $1K–7K/yr (triage fee + consumable margin) | Year 1–2 broadly — legitimacy signal, geographic coverage, routing function. |
| NeuroPulse Depot | Mail-in backstop for all cases local partners cannot handle | All tasks. T2 same-day loaner. eMMC specialist recovery. Fluxgate precision calibration. | Full NeuroPulse internal training | All equipment in-house | Highest margin per task — backstop role not primary channel | Day 1 — required before any device ships. |

# 7. Remaining Weaknesses + Open Risks

| Severity | Issue | Current status + mitigation |
| --- | --- | --- |
| CRITICAL | Zero published clinical trials — 35-trial gap vs Vielight | Every clinical adoption decision, press mention, and regulatory discussion will reference Vielight's 35+ RCTs against NeuroPulse's zero. SBIR Phase I + academic seeding strategy is correct but takes 2–3 years to produce published data. PRIORITY: Commission SBIR Phase I immediately. First contacts: Rashidi-Ranjbar (Toronto PBM MCI RCT), Jog (UCLA HD-tDCS), Naeser (VA Boston TBI/PTSD). The 33-entry bibliography supports 'consistent with established evidence on each component modality' framing at launch. |
| CRITICAL | 400 mW/cm² regulatory opinion not yet obtained | This number cannot appear in any public-facing material, investor deck, or product spec sheet until outside regulatory counsel confirms it is consistent with FDA-exempt wellness positioning and does not implicitly adopt Vielight's medical indications under FTC implied claim doctrine. Commission immediately: $8,000–15,000, 3–5 weeks. Until opinion is in hand, all materials must use '200 mW/cm²' or omit irradiance entirely. |
| CRITICAL | 'NeuroPulse' is an uncleared placeholder name | Trademark search and clearance in US, EU, Canada, Australia required before any external conversation — researcher outreach, investor pitch, crowdfunding design, regulatory filing. Budget $15,000–25,000. If name conflicts, better to know now than after any public exposure. |
| MODERATE | Scientific Advisory Board not formed | Device markets multi-modal 40Hz neuromodulation without any SAB credentialing. Priority: Li-Huei Tsai (MIT, 40Hz GENUS founder) as SAB member — NOT trial PI (Cognito conflict). Contact: 617-324-0305. Glen Jeffery (UCL, retinal PBM) and Margaret Naeser (VA Boston) also natural SAB candidates. 5-person SAB: $50,000–80,000/yr. Must be announced before or concurrent with clinical researcher outreach. |
| MODERATE | Vulnerable population edge cases in research consent | Depression/PTSD category patients who contribute data and later wish to withdraw may not understand that already-anonymised data cannot be removed from an ongoing study. Withdrawal flow must disclose at per-project consent: 'Once your anonymised data is included in a study, individual withdrawal is not possible — future contributions can always be stopped.' Required by research ethics (Common Rule) and must appear in consent document before IRB approval of any study using this platform. |
| MODERATE | 45W GaN charger not universally available | Full therapeutic PBM requires 45W USB-C PD. Users with only a 5V/2A phone charger get EEG-only mode. Decision: include branded 45W GaN charger in Home Standard box ($8–12 BOM). App clearly communicates 'power level: full therapeutic' vs 'power level: reduced' at session start. Eliminates friction at first use. |
| STRUCTURAL | Fluxgate calibration still requires service center visit | Reference photodiode eliminates PBM calibration visits. NTC cross-calibration and ADS1299 self-cal eliminate temperature and EEG drift. But fluxgate sensitivity drift (scale factor) requires a Helmholtz test coil providing a known field magnitude — not achievable via geomagnetic comparison alone. Residual requirement: Tier B service center every 3–5 years. Mitigated by: geomagnetic comparison detecting when drift is severe enough to schedule, fleet SHDR trend prediction giving non-urgent lead time, 3–5 year interval being long enough for Tier B network to be well-established first. |
| STRUCTURAL | iOS/Android OS updates can disrupt app function | Annual major iOS/Android releases can break USB accessory APIs, background execution, or health APIs without warning. Committed mitigation: participate in Apple/Google developer beta programmes, 7-day OS compatibility SLA, autonomous Mode 3 operation (pre-programmed session from power bank) as structural fallback when app is temporarily broken. This cannot be fully designed away — it is a dependency on platform vendors. |

# 8. Open Items + Completed Decisions

## 8.1 Pending decisions — action required before first tooling cut

- Product name trademark search and clearance (US/EU/CA/AU) — before ANY external conversation

- 400 mW/cm² regulatory opinion letter — before any marketing material references irradiance

- Li-Huei Tsai SAB outreach (617-324-0305) — SAB role only, not trial PI

- LED emitter pulse current rating verification (660nm + 808–830nm candidates) — confirm 120–180mA continuous FPC rating before component commitment

- Zone module FPC layout specification — 6mm pitch, lever-actuated ZIF, reference photodiode per zone, plasma-activated PDMS windows — must be in PCB layout before first cut

- Hub tooling: intranasal probe dock + anchor posts for all interface covers + large-radius Boa cable channel + tool-free fan access (quarter-turn captive fastener) — all in hub tooling before first cut

- Shell tooling: anchor posts for zone slot covers (5 positions, colour-coded) and accessory port covers (3 positions) — in headset shell tooling before first cut

- Lens tooling: sliding rail mount geometry + N42 magnet positions (≥1mm polymer wall all faces) + AgNW coating spec + hard coat over AgNW + EC lens driver electrical contacts spec — before lens tooling cut

- Goggle arm tooling: anchor hook for lens rim guard tether — before arm tooling cut

- Partner optician network engagement — contract for S3 Rx clip programme before device ships

- eMMC partition architecture + UHDR/SHDR separate encryption scheme — in firmware specification before first firmware is written

- Dual-bank OTA bootloader — must be in bootloader from first firmware line

- SBIR Phase I application — initiate immediately after company formation

- First researcher contact: Rashidi-Ranjbar (neda.rashidi-ranjbar@unityhealth.to), then Jog (mjog@mednet.ucla.edu), then Naeser (mnaeser@bu.edu)

## 8.2 Completed decisions — locked

- Two-tier product strategy (T1 consumer wellness / T2 clinical) with shared platform

- 600 LEDs (300×660nm + 300×808–830nm) FPC zone modules from launch — no Rev A/Rev B split

- 400 mW/cm² peak pulsed via firmware duty cycle (Path A, $0 hardware) — pending regulatory opinion

- 21-channel EEG: FC3, FC4, Oz, A1/A2 (on VNS clips) — full 10-20 + auricular references

- STM32G071 safety MCU (not G031 — SRAM error corrected) — +$0.45 BOM

- No NFC anywhere — optical code + resistive pogo authentication

- USB-C wired-first default, BT/Wi-Fi toggle, antennas in hub not headset

- Boa-style dial + spring-decoupled electrode pods — one adult SKU 52–62cm

- 5-layer passive EMF shielding + active Helmholtz cancellation + palladium fabric (permanent shielding claim)

- Auricular VNS clip = HRV PPG + A1/A2 EEG references (2 spare conductors in existing 6-pin cable)

- Standard lens (35% VLT, 108 micro-LEDs/lens, S1 included) + EC premium (+$89, 5–75% VLT) + S3 Rx clip (both lenses)

- Internal processor retained ($85 BOM) — autonomous closed-loop is primary competitive moat

- UHDR/SHDR terminology framework — complete separation, never linked by design

- Clinical consent engine: use case library → minimum necessary data mapper → consent document generator

- Research partnership: a priori consent (contact/category/blanket), results opt-in, suggestion portal, crowdfunding

- POA workflow with human document review, jurisdiction flagging, annual re-verification

- Service network: Tier A optical (primary) + Tier B electronics + Tier C retail + depot backstop

- All interface protection covers tethered to headset, anchor posts in tooling

- Predictive maintenance: SHDR fleet database → fleet-trained RUL model → deployed back to devices via OTA

- Design brief Revision 1, Revision 2, bibliography (33 entries), researcher candidates (12), and this Revision 3 generated as Word documents