NeuroPulse

Master Design Brief

Closed-Loop Multi-Modal Neuromodulation Platform — Full Design Review

Revised: May 3, 2026  ·  Revision 2 — incorporates all design decisions

Key decisions added since Revision 1 - Visual lens system: Inner micro-LED array (108 LEDs/lens, 660nm + 808–830nm) — not a display. Standard lens (35% VLT) + EC premium option (5–75% VLT, +$89). Three snap-on shade variants (S1 opaque included, S2 polarising $24, S3 prescription $49+Rx). Mode F: invisible NIR retinal walk daily protocol. - PBM irradiance: Upgraded to 400 mW/cm² peak pulsed via firmware-only (Path A, $0 BOM). 25% duty cycle enforced by existing NTC thermal interlock. Matches Vielight Neuro Pro 2 spec. Regulatory review required before marketing claim. - LED count: Upgraded to 600 LEDs total (300 × 660nm + 300 × 808–830nm per headset). FPC zone modules from launch. Eliminates Rev A/Rev B sequencing — ships with full 600-LED design from day one. +$52–58 BOM, ~25% power efficiency gain. - EEG: Upgraded to 21 channels (from 16). Added FC3, FC4 (motor cortex TMS targeting), Oz (photoparoxysmal detection, visual SSVEP), A1/A2 (linked-ear reference on VNS auricular clips). Full 10-20 + auricular references. +$15 BOM. Matches/exceeds CONTEC KT88-2400/3200 channel count with superior 24-bit/2kHz ADC. - Safety MCU: Confirmed STM32G071 (not G031). G031 has only 8KB SRAM — insufficient for EMF firmware tasks. G071: 36KB SRAM, same footprint/toolchain, +$0.45 BOM. - EMF shielding: 5-layer passive (CFRP + mu-metal + silver fabric + EMI foam + port filters) + active Helmholtz cancellation. 35–45dB ELF magnetic, 40–60dB RF. Three firmware additions required: TMS-gated cancellation, adaptive notch at BES/tACS stimulus freq, synchronous Helmholtz subtraction from EEG. Non-conductive CFRP window at TMS coil site. - Power: PBM = 85-90% of all power draw. Standard protocol: 20W (15V/2A, 45W charger). Tier 2 peak: 74W (100W EPR). 10,000mAh power bank ≈ 90min standard session. Supercapacitor buffer (10F/5V) recommended to absorb LED duty-cycle current transients. - Connectivity: USB-C default (zero RF). BT 5.3 + Wi-Fi 6 via single rear toggle. Antennas in control hub not headset. Phone-as-power supported (cable length = EMF mitigation, phone-in-pocket). No internal battery in headset. - Phone-tethered design: Analysed and rejected as primary configuration. App can kill stimulation (safety MCU handles safely). iOS requires MFi certification for background execution (post-launch). Processor stack ($85 BOM) retained — autonomous closed-loop is primary competitive moat. - Electrode consumables: Semi-dry hydrogel tips (snap-off pods) — 30–60 sessions, $12–16/8-pack. VNS clip pads — 20–40 sessions, $8/2-pack. Audio cup foam — 6–12 months, $24/set.

# 1. Executive Summary

NeuroPulse is a two-tier closed-loop neuromodulation platform — the first wearable brain device integrating photobiomodulation, EEG, transcranial electrical stimulation, vagus nerve stimulation, audio entrainment, and visual stimulation with autonomous mobile operation and multi-layer EMF shielding.

## 1.1 Product tiers

| NeuroPulse Home — Tier 1 Consumer Wellness · FDA-exempt · 8 modalities - 8 modalities: PBM transcranial (660+810nm), PBM intranasal bilateral, EEG 8-ch neurofeedback, BES/tACS entrainment, tDCS cortical priming, VNS+HRV auricular, audio entrainment, visual stimulation - Pricing: $449 (Core) · $599 (Lite) · $849 (Standard ★) · $1,199 (Premium) - Target: Health-conscious adults 35–65, biohackers, premium wellness buyers - Time to market: 12–18 months from design freeze - No mandatory subscription; full function on-device from day one | NeuroPulse Pro — Tier 2 Clinical/Research · FDA 510(k) target · All modalities - Adds: TMS focal figure-8 coil (rTMS/TBS), 21-ch wet gel qEEG + sLORETA, 1170nm deep PBM (35–40mm), clinical tACS (≤4mA, 16-ch), cervical VNS option, HIPAA cloud, EHR (FHIR R4), research scripting API - Pricing: $4,999 (Pro Entry) · $13,999 (Pro Full) + $1,800/yr service - Target: Neurology clinics (CPT 90867–90869 reimbursed TMS), psychiatry, TBI rehab, academic research - FDA pathway: 510(k) modular predicate — 18–36 months, $2–5M - Upgrade path: T1 users receive hardware add-on kit + Pro certificate |
| --- | --- |

## 1.2 Founding design principles

Shared platform (one production line, two markets) · Wired-first USB-C default (zero RF at scalp) · Autonomous closed-loop EEG-adaptive stimulation without phone · Multi-layer EMF shielding (only consumer brain wearable with measured shielding) · Modular field-upgradeability via snap-in zone modules · No mandatory subscription

# 2. Modality Stack — Full Specification

| Modality | Consumer name | Tier | Tier 1 specification | Tier 2 additions |
| --- | --- | --- | --- | --- |
| PBM transcranial | Photobiomodulation | T1+T2 | 660–670nm + 808–830nm LEDs · 5 independently addressable zones · 300 LEDs per wavelength (600 total) on FPC strips · 10–400 mW/cm² peak pulsed (≤25% duty cycle at 400 mW/cm²) · 10–200 mW/cm² CW · closed-loop photodiode feedback per zone · real-time J/cm² dose metering · 0–10,000 Hz pulse · thermal-adaptive duty cycle (NTC per zone) | Adds 1170nm laser diode deep channel · 35–40mm subcortical depth · TEC stabilisation · 10–1,000 mW/cm² at 1170nm · all 5 zones at all 3 wavelengths |
| PBM intranasal | Intranasal light therapy | T1+T2 | Bilateral Y-probe · 660nm + 808–830nm per probe · 15/20/25mm depth stops (tactile click rings) · photodiode dose metering + contact sensing · optical code + pogo pin resistive sleeve auth (zero NFC) · hygiene sleeve consumable (30-pack, $19/pack or $19/mo subscription) | Higher irradiance ceiling (150 mW/cm²) · EEG-adaptive theta protocol · formal clinical integration |
| EEG / qEEG | Brain monitoring | T1+T2 | 8-ch semi-dry hydrogel electrodes · Fp1/2, F3/4, C3/4, P3/4 · 500Hz · 24-bit ADC · delta→high gamma · real-time neurofeedback · power bands + coherence · EDF+ export · replaceable hydrogel tips (30–60 sessions, $12–16/8-pack) | 21-ch wet gel upgrade · Full 10-20 + FC3 + FC4 + Oz + A1/A2 (on VNS clips) · <1µV RMS noise · CMRR >120dB · sLORETA source imaging · ERP paradigms · LSL streaming · FHIR R4 EHR export · research scripting API |
| Brainwave entrainment stimulation (BES/tACS) | Brainwave entrainment | T1+T2 | 0.5–40Hz · ≤1mA · 3 electrode pairs · charge-balanced biphasic · per-electrode impedance monitoring · consumer parameter ranges · adaptive notch prevents active cancellation from cancelling stimulus | 0.5–100Hz · ≤4mA · 16-ch arbitrary waveform · clinician-controlled full parameter exposure |
| Cortical priming stimulation (tDCS) | Cortical priming | T1+T2 | 0.1–2mA DC · 5–30min sessions · 30s ramp up/down · ≤3 electrode pairs · 40µC/cm² limit hardware-enforced in safety MCU | 0.1–4mA · up to 8 simultaneous pairs · 10–60s programmable ramp · HD-tDCS montage library |
| VNS + HRV | Vagus nerve stimulation | T1+T2 | Auricular clip · auricular branch CN X · 1–25Hz · ≤2mA · biphasic charge-balanced · wellness params · PPG HRV (808–830nm) in same clip · force contact confirmation · PDMS hydrogel pad (20–40 session, $8/2-pack) · A1/A2 EEG references on clip contact pads | 0–5mA · clinical biphasic · optional cervical electrode accessory ($49) · full autonomic monitoring |
| Neural audio entrainment | Audio entrainment | T1+T2 | Over-ear planar magnetic 40mm · bone conduction at mastoid · binaural beats (0.5–40Hz) · isochronic tones · pink/brown noise · VNS clip cutout in lower rim · EEG-adaptive frequency · Helmholtz coil orientation optimised to reduce magnetic coupling to audio membrane | EEG-adaptive bilateral clinical EMDR audio mode · session guidance overlay |
| Visual neural entrainment | Visual stimulation | T1+T2 | 108 micro-LEDs per lens (660nm + 808–830nm) · 6 zones per eye (L, CL, CR, R, Top, Bot) · inner PDMS diffuser film · ITO outer EMF coating (ground to chassis) · IR eye-open detection (940nm) · photic driving 0.5–100Hz · EMDR L/R alternation · retinal PBM mode · Mode F (NIR retinal walk, invisible to user) · IEC 62471 hardware MPE limit · hall sensor goggle-lift cutoff | Clinical EMDR protocols · EEG-adaptive photic phase-locked to live oscillation · photoparoxysmal EEG detection → auto-halt <200ms at Oz |
| TMS | — (Tier 2 clinical) | T2 only | Not available in Tier 1 | Focal figure-8 coil · 0.1–0.5T · rTMS + TBS protocols · active coil cooling · cap-bank in separate hub · non-conductive CFRP window at coil site (prevents eddy currents) · TMS-gated EMF cancellation (safety MCU gates Helmholtz off 5ms before each pulse) |

## 2.1 PBM Frequency Preset Library

| Preset | Freq | Primary zones | Secondary modalities | Evidence + rationale |
| --- | --- | --- | --- | --- |
| Gamma clarity | 40 Hz | Prefrontal + temporal | BES gamma · audio 40Hz binaural | MIT Tsai: 40Hz reduces amyloid-β and tau (Nature 2016). Strongest transcranial PBM evidence. Gamma entrains memory binding. Phase 2A RCT (n=15 AD): reduced hippocampal atrophy, better face-name recall. Primary clinical AD protocol. |
| Alpha calm | 10 Hz | Parietal + occipital + DMN | VNS · HRV coherence · alpha BES | Most-published transcranial PBM frequency. Vielight's entire 35+ trial track record at 10Hz. Drives thalamocortical alpha loops. Measurably increases resting EEG alpha power post-session. |
| Theta memory | 6 Hz | Temporal + intranasal bilateral | Intranasal bilateral · theta BES | Hippocampal dominant rhythm for memory encoding. Intranasal delivers theta-coupled light to olfactory bulb → hippocampus via direct pathway. Age-related cognitive decline and MCI protocol. |
| Sleep deep | 2 Hz | Full cap + intranasal | Delta audio · VNS parasympathetic | Delta supports slow-wave sleep architecture and glymphatic waste clearance. Pre-sleep use. Outdoor/environment flicker most problematic at this frequency — full shade recommended. |
| Gamma + theta coupled | 40 + 6 Hz | Transcranial 40Hz / intranasal 6Hz | Audio theta-gamma coupling | Mirrors hippocampal cross-frequency coupling architecture (40Hz gamma nested in 6Hz theta). No existing device implements simultaneous dual-frequency split-zone protocol. NeuroPulse unique. |
| Focus prime | 20 Hz | Prefrontal bilateral | tDCS cortical priming | Beta drives active cognitive engagement. tDCS anodal priming amplifies motor/cognitive cortex excitability. Pre-task executive function protocol. |
| Vascular baseline | CW (0 Hz) | Intranasal only | None | Continuous wave maximises nasal capillary blood irradiation for systemic PBM. Research baseline condition and systemic wellness protocol. |

# 3. Hardware Specifications

| System | Specification | Key decisions + rationale |
| --- | --- | --- |
| Main processor | NXP i.MX RT1062 · Cortex-M7 · 600MHz · FPU+DSP+SIMD · 1MB on-chip SRAM + 32MB LPSDR4 SDRAM · 8GB eMMC 5.1 · USB-HS OTG (480 Mbps) · FreeRTOS · ~1.1% CPU utilisation at full load including all EMF firmware tasks | 98.9% CPU headroom reserved for future on-device ML (TF Lite Micro with CMSIS-NN). Rich peripheral set (FlexPWM, multi-SPI, USB-HS OTG) replaces external ICs. Autonomous closed-loop: FFT band-power at 10Hz, state classification, parameter update all run in headset without connected device. |
| Safety MCU | STM32G071 · Cortex-M0+ · 64MHz · 36KB SRAM · 128KB flash · bare-metal · owns all stimulation GPIO enable lines · 1.5s watchdog · SPI heartbeat from main processor · Q15 fixed-point CMSIS-DSP (no FPU) | UPGRADED from G031 (8KB SRAM insufficient for EMF firmware tasks — design error caught in timing analysis). G071 drop-in replacement at +$0.45 BOM. IEC 62304 Class C scope: ~500 lines bare-metal code, formally verifiable. Hardware key: app crash cannot cause unsafe stimulation. All-stimulation cutoff <50ms from any fault trigger. |
| PBM LED array | 600 total LEDs per headset · 300 × 660–670nm (red CCO peak) + 300 × 808–830nm (NIR CCO peak) · 5 zones × 60 LEDs per wavelength per zone · FPC (flexible PCB) strips per zone · 6mm inter-LED centre-to-centre pitch · drive current 120–180mA per LED (vs 700–1000mA with 100 LEDs) | Launches at 600 LEDs from day one (no Rev A/Rev B split). FPC enables scalp-conforming deployment. At 120–180mA/LED: near-uniform field (±15–25% variation vs ±300–400% at 100 LEDs), LED L70 lifetime 80,000–100,000 hrs, wall-plug efficiency 38–46% (vs 28–33% at high drive). 25% reduction in total PBM power draw for same optical dose. |
| PBM irradiance | 10–400 mW/cm² peak pulsed · 10–200 mW/cm² CW · thermal-adaptive duty cycle (hardware NTC throttle per zone) · real-time J/cm² dose metering via closed-loop photodiode | 400 mW/cm² peak pulsed achieved via firmware-only (Path A, $0 BOM). IEC 60601 scalp temperature limit (42°C) caps CW at ~200 mW/cm²; at ≤25% duty cycle, 400 mW/cm² peak stays within thermal limit. Matches Vielight Neuro Pro 2 headline specification. Regulatory review required before marketing claim. NeuroPulse uniquely publishes both peak irradiance AND real-time J/cm² dose — Vielight does not. |
| EEG system (T1) | 8-ch semi-dry hydrogel · Fp1/2, F3/4, C3/4, P3/4 · spring-loaded pods (80–120g, ±12mm travel) · 500Hz · 24-bit ADC · <2µV peak-to-peak noise · CMRR ≥100dB · DRL-driven shell bonding (shell = active EEG shield) | Spring-decoupled electrode pods: 3–4g/mm spring maintains 80–120g contact force independent of dial tightness — resolves fundamental comfort-vs-signal conflict of all existing consumer EEG headsets. Semi-dry hydrogel: no prep, <5min setup. Replaceable tips: snap-off bayonet (30–60 sessions, $12–16/8-pack). |
| EEG system (T2) | 21-ch wet gel cap · Full 10-20 system + FC3 + FC4 + Oz + A1 + A2 · <1µV RMS noise · CMRR >120dB · input impedance ≥100MΩ · 250–2,000Hz selectable · 3× ADS1299 (8-ch ICs) · sLORETA source imaging · EDF+ · LSL · FHIR R4 | 21-ch justification: FC3/FC4 = M1 motor cortex targeting for TMS (required for TMS targeting verification); Oz = maximal photoparoxysmal response site (enhances seizure safety interlock); A1/A2 on VNS clips = linked-ear reference for normative database comparison (NeuroGuide, LORETA Key). Exceeds CONTEC KT88-2400 (19-ch) and KT88-3200 (32-ch) on every ADC parameter (12-bit vs 24-bit, 200Hz vs 2kHz, ≥90dB CMRR vs ≥120dB). |
| Fit system | Boa-style occipital dial (10cm range, 0.5mm/click, 50,000-cycle rated) · 5-position forehead bridge (5mm steps) · spring-loaded electrode pods (±12mm, Shore 30A silicone integrated grip) · temporal stability wings (snap-on, stored in hub) · 1 adult SKU covers 52–62cm | 1 SKU covers 90%+ of adults. No velcro (degrades, loses pads). Autoclavable arch assembly (quarter-turn field-replaceable) for T2 clinical use — $39/pair replacement. Non-removable integrated silicone grip is autoclavable, eliminates loose parts in clinical settings. |
| EMF shielding | Layer 1: CFRP outer (30–50dB RF) · Layer 2: 0.2mm mu-metal liner (15–25dB ELF magnetic) · Layer 3: Silver-coated polyester inner liner (40–60dB RF) · Layer 4: Carbon-loaded EMI foam (cavity resonance suppression) · Layer 5: USB-C + accessory port filters (30–50dB) · Layer 6: Active Helmholtz cancellation (3-axis fluxgate + coil pairs, adds 20–25dB ELF magnetic) · All shell layers bonded to EEG DRL output | Only consumer brain wearable with measured multi-layer EMF shielding. Combined: 35–45dB ELF magnetic, 40–60dB RF. +$119 BOM. Non-conductive CFRP window at TMS site prevents eddy current braking (20–40% TMS field loss avoided). TMS-gated cancellation, adaptive notch at BES/tACS stimulus freq, and synchronous Helmholtz-EEG subtraction are three firmware additions required — all use information already present in the safety MCU's session schedule. DRL bonding is standard shielded-EEG practice, eliminates common-mode coupling. |
| Connectivity | USB-C 3.2 Gen1 default (zero RF, <1ms latency, PD power delivery) · BT 5.3 LE Audio (10mW) · Wi-Fi 6 (2.4/5GHz) · Single rear toggle (press=cycle, hold=pair, double=USB-C) · LED ring per mode · Antennas in control hub, not headset · MagSafe-style breakaway at shoulder · 90° connector exits rearward from headset | Hub-mounted antennas reduce scalp RF 20–30dB vs headset-mounted even in wireless mode. Phone-as-power: 1.5m cable = phone-in-pocket at sufficient distance that cellular RF falls to <0.05µW/cm² at scalp (ambient background level). Headset processor retained despite $85 BOM cost — autonomous closed-loop is primary competitive moat; removing processor saves 10% COGS while permanently destroying it. |
| Operating modes | Mode 1 Connected: real-time streaming <1ms · Mode 2 Programming: app uploads session protocol <5s · Mode 3 Autonomous: pre-programmed, runs from any USB-C PD power bank, closed-loop EEG-adaptive included · Mode 4 Download: reconnect USB-C, app retrieves EDF+ + parameter logs | 8GB eMMC: 1,000+ T1 sessions (30min, 8-ch compressed ≈7MB each) or 60 T2 sessions (60min, 21-ch compressed ≈120MB each). Closed-loop engine runs fully autonomously in Mode 3 — no competitor offers this. Session protocol format: binary, cryptographically signed by app before headset accepts it. |
| Status indicators | Left temple: green power LED (breathes slowly at idle) · Right temple: amber in-use LED (pulse rate mirrors session frequency as real-time confirmation) · Fault: power LED red blink · Stealth mode: app-controlled suppress (safety fault always fires) · Both forward-facing, visible at eye level | Amber pulse rate as session frequency confirmation is unique in industry — clinical caregiver confirms correct protocol from across room without touching app or patient. Spatially distinct LEDs require no learned color code: left=power always, right=activity always. |
| Visual lens system | Standard lens: 35% VLT polycarbonate shield · 108 micro-LEDs/lens (660nm + 808–830nm) · 6 zones/eye · PDMS inner diffuser film (75% haze, 15% in optical zone) · ITO outer conductive coating (EMF ground) · IR proximity sensors (940nm) for eye-open detection · Hall sensor: goggle lift = LED cutoff · S1 opaque shade included | Not a display — a controllable light field. $18–28 BOM per lens vs $180–350 for waveguide display, with identical functional capability for photic driving, EMDR, and retinal PBM. Mode F (NIR retinal walk): 808–830nm is invisible, user appears to wear normal sport glasses during ongoing low-dose retinal PBM. No waveguide bulk or camera array needed. |
| EC lens (premium) | Bistable electrochromic film · 5–75% VLT · app + button controlled · 2s transition · ~180mW during transition · ~15mW hold (bistable holds state without power) · always clears to 75% on power restore (safety failsafe) | +$89 upgrade / $129 standalone replacement. Eliminates S2 polarising clip need. Bistable: only 0.45Wh additional per 30-min session — negligible runtime impact. Retained S1 opaque shade in box: EC transition (2s) too slow for some clinical switching; S1 is instant. Both EC and standard lenses accept S3 prescription clip. |
| Snap-on shade system | S1 opaque (<0.5% VLT, included) · S2 polarising (~12% VLT, $24, standard lens only) · S3 prescription (1.67 high-index corrective lens, $49 carrier + $49–139 Rx insert) · Magnetic snap: 6× N52 neodymium inserts in lens rim · Pull-off force: ~1.2kg (secure during light exercise, removable one-handed) | S3 prescription clip is accessibility-critical: ~75% of adults 40+ need vision correction; ~55–70% of target demographic. Without S3, majority of target market excluded from open-eye session modes (EMDR, Mode F NIR walk). S3 compatible with both standard and EC lens. Rx insert renews every 12–24 months = recurring revenue with zero marginal marketing cost. App-based Rx submission → partner optician network → 5–7 day fulfilment. |

# 4. Configurations + Pricing

## 4.1 Integrated system configurations

| Configuration | BOM est. | Mfg. COGS | Retail price | Gross margin | Included modalities | Target customer |
| --- | --- | --- | --- | --- | --- | --- |
| Core — EEG only | $165 | $255 | $449 | 43% | 4-ch EEG · USB-C · BT · Wi-Fi · EMF shielding · full processor stack · 8GB storage · indicators | Cautious entry buyer; neurofeedback researchers; B2B modular base unit |
| Home Lite | $260 | $365 | $599 | 39% | All Core + PBM 5 zones (660+810nm, 600 LEDs) · 8-ch EEG · VNS+HRV clip | Wellness-focused consumer seeking Neuronic/Vielight alternative at lower price |
| Home Standard ★ | $395 | $530 | $849 | 38% | All T1 modalities: PBM · intranasal · 8-ch EEG · BES · tDCS · VNS+HRV · audio cups · standard visual lens · S1 shade included | Primary flagship; biohacker; health-conscious adult 35–65; primary marketing SKU |
| Home Premium | $450 | $610 | $1,199 | 49% | All T1 modalities + EC lens upgrade (+$89) + extended 2yr warranty + priority support | High-income buyer preferring EC lens from day one; premium experience segment |
| Pro Entry | $820 | $1,350 | $4,999 | 73% | All T1 + 21-ch qEEG wet gel · 1170nm deep PBM · clinical tACS · HIPAA cloud · EDF+ · sLORETA | Neurology clinic; concussion rehab centre; academic research PI (IRB exemption seeding) |
| Pro Full | $1,480 | $2,600 | $13,999 | 81% | All T2: Pro Entry + TMS hub · multi-patient dashboard · research mode · LSL streaming · scripting API · FHIR R4 · $1,800/yr service | Academic medical centre; psychiatry clinic; neuromodulation research lab |

## 4.2 Power requirements by operating mode

PBM is 85–90% of all power draw in every active mode. All other modalities combined account for 2–5W. The 25% LED efficiency improvement from 600-LED design reduces standard session draw from ~20W to ~16–17W, extending power bank runtime by ~25% vs prior specification.

| Operating mode | Total draw | Min USB-C PD | Runtime (10,000mAh) | Notes |
| --- | --- | --- | --- | --- |
| Standby / idle | 1W | 5V/0.5A | ~330 hours | Any USB source including phone USB-A |
| EEG monitoring only | 2.5W | 5V/1A | ~130 hours | Laptop USB port sufficient. Full closed-loop neurofeedback operational. |
| Standard T1 protocol ★ | ~17–20W | 15V/2A (45W) | ~95–110 min | 5 zones 660+810nm + all modalities. 45W MacBook charger = everyday supply. (Lower end from 600-LED efficiency gain.) |
| All modalities T1 peak | ~45–50W | 20V/3A (65W) | ~38–42 min | Maximum irradiance all zones simultaneously. |
| T2 standard (+ 1170nm) | ~44–46W | 20V/3A (65W) | ~41–43 min | Typical clinical session with deep PBM. |
| T2 peak (no TMS) | ~70–74W | 20V/5A (100W EPR) | ~24–27 min | All 3 wavelengths at maximum. |
| TMS hub (T2 add-on) | 30–50W | Dedicated supply | Wall recommended | Separate USB-C 100W or barrel-jack. Not from headset PD bus. 10F/5V supercapacitor buffer in hub recommended to absorb duty-cycle transients. |

## 4.3 Modular pricing (field upgrades and clinical procurement)

Integrated systems are the primary sales motion. Modular pricing exists for field upgrades (existing users adding capability), clinical line-item procurement accounting, and replacement components. Integrated systems save 38–83% vs modular assembly.

| Module | BOM est. | Modular retail | Integrated value | Notes |
| --- | --- | --- | --- | --- |
| PBM zones — 5-zone 660+810nm (600 LEDs, FPC) | $155–170 | $349 | $255 | Largest single BOM item. 600-LED FPC adds ~$52–58 vs prior 100-LED spec. Drives 85% of power draw. |
| EEG 8-channel expansion (semi-dry hydrogel pods) | $45 | $129 | $90 | Same snap-in zone slot system. Includes one set of hydrogel tips. |
| BES + tDCS combined stimulator | $35 | $149 | $100 | Shared electrode array with EEG via fast analog switch matrix. |
| VNS auricular clip + HRV + A1/A2 references | $26 | $99 | $65 | Magnetic 6-pin accessory port. PPG + VNS + EEG reference in single unit. |
| Intranasal PBM applicator | $35 | $149 | $90 | Bilateral, 2λ, 3 depth stops. Sleeve subscription ($19/mo) = recurring revenue. |
| Over-ear audio cups | $55 | $179 | $130 | Bayonet quarter-turn. Conductive mesh ear face (40dB RF, acoustically transparent). |
| Visual goggle shield (standard lens) | $42 | $149 | $100 | 108 micro-LEDs/lens. ITO EMF coating. S1 shade included. +$89 for EC upgrade. |
| EC lens upgrade | $47–67 | $89 upgrade / $129 standalone | $130 | Bistable electrochromic. Replaces S2 clip need. Both lens options accept S3 prescription clip. |
| EEG 21-ch pro wet gel expansion (T2) | $80 | $279 | $200 | Includes FC3/FC4/Oz pod set + A1/A2 clip contacts. Tier 2 only. |
| 1170nm deep PBM — all 5 zones (T2) | $250 | $899 | $700 | Laser diode + TEC. 35–40mm subcortical depth. Tier 2 only. |
| TMS hub (T2, separate unit) | $200 | $2,999 | $2,400 | Separate chassis. Figure-8 coil. USB-C to headset. Supercapacitor buffer for pulse transients. |

# 5. Accessory Catalogue + Consumables

| Accessory | Price | Compatible | Specification | Market rationale |
| --- | --- | --- | --- | --- |
| EC lens upgrade | $89 upgrade / $129 standalone | T1 + T2 | Bistable electrochromic 5–75% VLT · 2s transition · ~15mW hold · clears to 75% on power restore · same micro-LED array as standard · both S1 and S3 accessories compatible | Eliminates S2 polarising clip for EC users. App-controlled opacity. Premium lens experience. Negligible power impact (0.45Wh/30-min session). |
| S1 opaque shade | Included / $18 replacement | Std + EC lens | <0.5% VLT · 6× N52 magnetic snap (insert-moulded in lens rim) · foam light seal · doubles as transport protection · instant (vs 2s EC transition) | Full immersion maximum stimulation fidelity. Retained with EC lens for protocols requiring instant mode change. Pull-off force ~1.2kg. |
| S2 polarising clip | $24 | Standard lens only | ~12% VLT · polarised polycarbonate · same magnetic snap · outdoor photic driving use | Not needed with EC lens (EC achieves same at 12% VLT via app). For standard lens users: outdoor conditions where sunlight (50,000+ lux) overwhelms LED stimulus at 35% base VLT. |
| S3 prescription clip | $49 carrier + $49–139 Rx insert | Std + EC lens | 1.67 high-index corrective lens · 14mm vertex distance · single-vision or progressive · same 6× N52 magnetic snap · app-based Rx submission → partner optician · 5–7 day fulfilment | ~75% of adults 40+ need vision correction. Without S3, majority of target demographic excluded from open-eye session modes (EMDR, Mode F NIR walk). Prescriptions renew 12–24 months = zero-marketing-cost recurring revenue. Available for both lens options. |
| Intranasal PBM applicator | $149 + $19/mo sleeves | T1 + T2 | Bilateral Y-probe · 660nm + 808–830nm per probe · 15/20/25mm depth stops · photodiode contact + dose sensing · optical code + pogo pin sleeve auth (no NFC, no EMF) · 30-pack hygiene sleeves included | Olfactory-hippocampal PBM pathway. Vielight's primary clinical differentiator — NeuroPulse matches + exceeds with bilateral design, dual wavelength, 3 depth stops. Sleeve subscription: 68–79% gross margin (consumable cost $4–6, retail $19). |
| VNS auricular clip + HRV | Included T1 Std+ / $89 module | T1 + T2 | Auricular electrical VNS · PPG HRV (808–830nm earlobe optical) · force contact sensor · PDMS hydrogel pads (20–40 sessions, $8/2-pack) · A1/A2 EEG reference contacts (2 spare conductors in existing 6-pin cable) | VNS current warms earlobe → improves PPG SNR. A1/A2 on clips: linked-ear EEG reference required for normative database comparison — 'free' capability using existing clip hardware. |
| Over-ear audio cups | Included T1 Std+ / $89 module | T1 + T2 | 40mm planar magnetic · bone conduction mastoid · VNS clip clearance in lower rim · conductive fabric mesh on ear face (40dB RF, 0.5mm spacing, acoustically transparent) · mild HF EQ correction (+2–3dB >8kHz for mesh roll-off) | Binaural beats require stereo isolation. Bone conduction adds dual-pathway entrainment. MIT: combined audio+visual 40Hz extends amyloid reduction to hippocampus. |
| Hygiene sleeve 30-pack (intranasal) | $19/pack · $19/mo subscription | T1 + T2 | Medical-grade silicone · laser-etched optical auth code · conductive pogo collar · single-use · zero NFC/EMF · wipe-clean seating surface | Only consumable in ecosystem. Clinical settings require documented single-use. Subscription = predictable MRR. Optical + resistive auth: zero EMF, <1ms read time, lower BOM per sleeve than NFC. |
| Electrode tip 8-pack (semi-dry hydrogel) | $12–16/pack · $9.99/mo subscription | T1 + T2 | Snap-off bayonet fitting · same geometry for scalp pods · 30–60 session lifespan · distilled water rinse and air-dry extends life · impedance degradation detected by EEG app | Primary T1 consumable after sleeves. App prompts replacement when impedance consistently elevated. Standardised bayonet shared with VNS pad where anatomically possible. |
| VNS clip contact pads (2-pack) | $8/pack | T1 + T2 | PDMS hydrogel · same snap-off design as electrode tips · 20–40 sessions (electrochemical degradation from VNS current shortens life vs scalp tips) | Second consumable. Combined electrode tip + sleeve subscriptions = $20–29/mo for daily users. |
| Audio cup foam (replacement set) | $24 | T1 + T2 | Memory foam over-ear pads · 6–12 month lifespan at daily use · hygiene replacement | Low-frequency consumable. Order-on-demand, not subscription. |
| Autoclavable arch set | $39/pair | T2 clinical | Shore 30A silicone forehead + occipital arches · quarter-turn field-replaceable · autoclave-rated · two-arch rotation for back-to-back clinical sessions | JCAHO and IPC infection control: velcro (competitors) cannot be autoclaved. Clinical billing requires sterilisable patient-contact surfaces. |
| Cervical VNS electrode | $49 | T2 only | Over carotid triangle CN X trunk · gel contacts · same magnetic 6-pin connector as auricular clip · swappable head | Some clinical VNS protocols require cervical delivery. gammaCore cleared predicates support regulatory path. Auricular standard for both tiers; cervical is clinical option. |
| Temporal stability wings | Included (stored in hub) | T1 + T2 | Curved silicone over-ear retention hooks · snap to lateral frame arms · sport and walking protocols · stored in dedicated hub slot when not in use | Mobile use without audio cups. Cannot be lost without losing control hub. Replaces velcro retention bands used by competitors. |

# 6. Competitive Positioning

| Feature | NeuroPulse Home | NeuroPulse Pro | Vielight Neuro Pro 2 | Neuronic Neuradiant 1070 | Sens.ai |
| --- | --- | --- | --- | --- | --- |
| Price | $849 | $4,999–13,999 | ~$5,000 | $3,000–4,995 | ~$1,500–2,000 + sub |
| PBM wavelengths | 660+810nm dual (2λ) | 660+810+1170nm (3λ) | 810nm only (1λ) | 1070nm only (1λ) | ~810nm (1λ) |
| Total LED count | 600 (300/wavelength) | 600 + 1170nm LDs | ~10 transcranial | 256–300 (1 wavelength) | ~7 midline only |
| PBM surface irradiance | 400 mW/cm² peak pulsed | 400 + 1,000 mW/cm² (1170nm) | 400 mW/cm² | Not specified | Not specified |
| Dose tracking (J/cm²) | Real-time per zone | Real-time per zone | No | No | No |
| Intranasal PBM | Bilateral, 2λ, 3 depths | Bilateral, 2λ, 3 depths | Yes (fixed single depth) | No | No |
| EEG channels | 8 semi-dry (24-bit) | 21 wet gel (24-bit, 2kHz) | None (roadmap) | None | 3 dry midline |
| Closed-loop EEG→stim | Yes — autonomous | Yes — all modalities | No | No | EEG→PBM freq only |
| BES/tACS | Yes (consumer name) | Yes, clinical grade | No | No | No |
| tDCS | Yes | Yes + HD-tDCS | No | No | No |
| TMS | No | Yes (focal figure-8) | No | No | No |
| VNS + HRV | Auricular electrical + PPG | + cervical option | Optical VNS (separate) | No | HRV only (no VNS) |
| Audio entrainment | Binaural + bone conduction | + clinical EMDR | No | No | No |
| Visual / photic | 108 LEDs/lens + EMDR | + EEG-adaptive, seizure detection | No | No | No |
| EMF shielding | 5-layer passive + active | 5-layer passive + active | None | None | None |
| Autonomous mobile | Pre-program + power bank | Pre-program + power bank | BT only (app required) | Portable, BT only | BT only |
| No mandatory subscription | Yes — full features | Yes + optional extras | Yes | Paywall on PLUS | Required $99–199/yr |
| Prescription lens support | S3 clip (both lenses) | S3 clip (both lenses) | No | No | No |
| Published clinical trials | None (new product) | None (new product) | 35+ RCTs | Limited | Ongoing |

# 7. Regulatory + Go-to-Market Strategy

## 7.1 Tier 1 — FDA-exempt wellness pathway

Positioning: general wellness device (same category as Muse, sens.ai, Apollo Neuro). No medical claims in consumer marketing. 'Brainwave entrainment stimulation' (not tACS); 'Cortical priming stimulation' (not tDCS). Avoids regulated clinical terminology while accurately describing mechanism. Time-to-market: 12–18 months from design freeze.

Tier 1 regulatory constraints:

- 1170nm laser diode absent from Tier 1 — removes Class 3B laser classification requirement

- 400 mW/cm² claim requires regulatory opinion letter before public marketing ($8,000–15,000 outside counsel) — get this before any material is published

- IEC 62471 photobiological safety hardware limits prevent any app-level override of retinal MPE ceiling

- Epilepsy interlock: EEG photoparoxysmal detection + Oz channel → goggle halt <200ms — unique safety feature no competitor has

## 7.2 Tier 2 — FDA 510(k) pathway

Modular predicate filing: TMS (NeuroStar/BrainsWay), tACS (Soterix/Neuroelectrics Starstim), transcutaneous VNS (electroCore gammaCore). Timeline: 18–36 months post-Tier 1 launch. Budget: $2–5M. CPT 90867–90869 TMS reimbursement enables immediate clinical revenue.

## 7.3 Go-to-market sequencing

| Phase | Activity |
| --- | --- |
| Year 1–2 | Tier 1 consumer launch. Target Neuronic, Vielight, sens.ai customers. Lead on dual wavelength + dose transparency + subscription-free + LED count + EMF shielding. Generate user base and recurring consumable revenue (sleeves, electrode tips). |
| Year 2–3 | Seed Tier 2 units into 3–5 research institutions under IRB exemptions (device donation + SBIR Phase I funding). Generate published clinical data — this is the primary strategy for closing the 35-trial gap vs Vielight. |
| Year 3–4 | Tier 2 commercial launch post-510(k) clearance. Sales targets: neurology (TMS), psychiatry, TBI rehabilitation, academic research institutions. HIPAA cloud platform + $1,800/yr service contract as recurring revenue engine. |
| Ongoing | Recurring revenue streams: intranasal sleeve subscription ($19/mo) · electrode tip subscription ($9.99/mo) · S3 prescription clip renewals ($49–139 every 12–24mo) · T2 service contracts ($1,800/yr) · optional advanced protocol library ($9.99/mo) · clinical data licensing to pharmaceutical partners |

# 8. Safety Architecture

## 8.1 Dual-processor isolation

- STM32G071 safety MCU runs bare-metal, completely independent of main processor — physically owns all stimulation enable GPIO lines

- Main processor sends SPI heartbeat every 200ms — 3 missed heartbeats (1.5s) triggers hardware cutoff of all stimulation within 50ms

- IEC 62304: Class C certification scope = ~500 lines bare-metal safety code. Main processor runs Class B separately

- App crash or OS kill cannot cause unsafe stimulation — hardware key is in safety MCU, not software

## 8.2 Modality-specific safety interlocks

| Modality | Interlock | Implementation |
| --- | --- | --- |
| EEG + Visual (photic) | Photoparoxysmal response detection → goggle halt | Oz electrode (newly added) as primary detection site. Pattern matched within 200ms (before clinical seizure event). Clinician-unlock required for 3–30Hz photic in T2. App screens PHx photosensitive epilepsy on first goggle use. |
| BES / tDCS | 40µC/cm² charge density hardware limit | Enforced in safety MCU — app cannot exceed regardless of software state. Hardware enforces even if main processor crashes. |
| Visual / retinal PBM | IEC 62471 retinal MPE hardware limit at 50% of exempt threshold | IR proximity sensors (940nm) detect eye-open and auto-reduce irradiance 20× in hardware (not software). Hall sensor: goggle lift = instant LED cutoff. Three independent hardware protection layers. |
| PBM (scalp thermal) | IEC 60601 42°C scalp limit | NTC thermistor per zone — hardware current throttle at 62°C LED junction. Same hardware enforces irradiance limit AND thermal safety simultaneously. Dual function. |
| TMS | Coil saturation + magnetometer protection | EMF cancellation gated off 5ms before each TMS pulse (safety MCU owns timing). 50ms post-pulse hold. Non-conductive CFRP window prevents eddy currents (field loss avoided). |
| VNS | Impedance check before each session | Safety MCU reads electrode impedance. Holds stimulation if contacts not confirmed. Biphasic charge-balanced waveform — net DC charge ≈0. |
| tDCS | DC path hardware-separated from AC | Separate circuit from BES. No interaction possible. 30s hardware-enforced ramp on/off prevents stimulus artefact. |
| All modalities | Firmware anti-fragility: CSPRNG session protocol signing | Session protocol cryptographically signed by app — headset rejects unsigned or corrupted protocols before execution. Prevents malformed third-party sessions from running. |

# 9. Open Items + Completed Decisions

## 9.1 Pending decisions

- Product family name: 'NeuroPulse' is a working name — branding decision required before any external conversations or marketing material

- Regulatory counsel engagement: 400 mW/cm² peak pulsed claim review must happen before this spec appears in any public-facing material ($8,000–15,000, outside counsel familiar with FDA PBM enforcement posture)

- LED emitter pulse current rating verification: confirm current 660nm and 808–830nm emitter candidates are rated for 120–180mA continuous in FPC format — verify with manufacturer before component commitment

- PCB tooling: Zone module FPC layouts must be finalized before first tooling cut. 21-ch EEG cap positions and A1/A2 clip contacts must be specified before VNS clip tooling

- Partner optician network: S3 prescription clip programme requires optician network agreement before product launch

- MFi certification planning: Required if iOS background execution needed for phone-tethered mode (not on critical path for launch). Timeline: 3–6 months, $10,000–25,000

- Scientific Advisory Board composition: Li-Huei Tsai (MIT) is priority first contact for 40Hz GENUS scientific credibility. Outreach should precede any public announcement about visual stimulation protocols

- Clinical trial partner prioritization: First contact order — Rashidi-Ranjbar (Toronto, PBM MCI) → Jog (UCLA, HD-tDCS) → George (MUSC, TMS infrastructure) → Naeser (VA Boston, TBI/PTSD cohort)

- Supercapacitor buffer specification: 10F/5V in control hub for PBM duty-cycle transient absorption — value confirmed by analysis, component selection not yet committed

## 9.2 Completed decisions (locked)

- Two-tier product strategy (Tier 1 consumer wellness / Tier 2 clinical)

- 8 Tier 1 modalities + TMS + 1170nm deep PBM + 21-ch qEEG for Tier 2

- Wavelengths: 660–670nm + 808–830nm (updated from 650+850nm) + 1170nm Tier 2

- 600-LED FPC zone modules from launch (no Rev A/Rev B sequencing) — $52–58 BOM premium justified by: 6× LED lifetime, 25% power efficiency gain, uniform field, competitive LED count superiority

- 400 mW/cm² peak pulsed via firmware duty cycle (Path A, $0 BOM additional hardware)

- 21-channel EEG montage: FC3, FC4, Oz, A1/A2 added to base 16-channel design — +$15 BOM, A1/A2 on existing VNS auricular clips

- STM32G071 safety MCU (not G031 — SRAM error corrected) — +$0.45 BOM

- No NFC anywhere in device: optical code + resistive pogo authentication for sleeve and accessories

- USB-C wired-first default with BT/Wi-Fi toggle, antennas in control hub not headset

- Boa-style dial fit system with spring-decoupled electrode pods

- 5-layer passive + active Helmholtz EMF shielding with three firmware fixes (TMS gating, BES notch, Helmholtz-EEG subtraction) and shell bonded to EEG DRL output

- Auricular VNS clip doubles as HRV PPG sensor + A1/A2 EEG references — two conductors in existing 6-pin cable

- Standard lens (35% VLT, micro-LED array, S1 included) + EC lens premium (+$89, 5–75% VLT bistable) both accepting S3 prescription clip

- Phone-tethered design analysed and retained as optional configuration, not primary. Internal processor retained ($85 BOM) — autonomous closed-loop is primary competitive moat

- Electrode consumables defined: hydrogel tips (30–60 sessions), VNS pads (20–40 sessions), audio foam (6–12 months)

- Design brief Revision 1, bibliography (33 entries), and researcher candidate list (12 researchers) generated as Word documents