NeuroPulse

Zone Module FPC — Procurement Requirements

Document ID:  NP-PROC-FPC-001

Revision:  A  —  2026-05-06

Status:  Active — must be reviewed before first purchase orders are placed

Applies to:  All components for Zone Module FPC (NP-HW-FPC-001 Rev C)

Owner:  Supply Chain / Hardware Engineering

Risk register:  NP-HW-FPC-001 Risk Register — items RISK-01, RISK-07, RISK-08 addressed here

This document defines mandatory procurement specifications that must appear on purchase orders, supplier qualification letters, and incoming inspection instructions. Specifications not stated here that conflict with supplier defaults will result in non-conforming parts being delivered without notification.

1.  SCOPE AND PURPOSE

This document specifies procurement requirements for components used in the NeuroPulse Zone Module Flexible Printed Circuit (NP-HW-FPC-001). It exists because several component parameters critical to device performance and safety are not enforced by default supplier processes — they must be explicitly stated on purchase orders or they will not be controlled. Requirements are derived from risk analysis in the companion risk register (NP-HW-FPC-001 Risk Register Rev A).

All specifications in this document are MANDATORY unless marked ADVISORY. A purchase order that omits a mandatory specification is non-conforming and must not be placed without written approval from the Hardware Engineering Lead.

| Component category | Risk register origin | Section |
| --- | --- | --- |
| LED emitters — 660 nm and 808–830 nm | RISK-08 (Vf binning) | §2 |
| LED driver IC — Infineon BCR421W | RISK-07 (variant confusion) | §3 |
| Zone module connector — Hirose FH34S / JAE FF03 | RISK-01 (wrong connector family) | §4 |
| FPC fabrication — copper type and weight | RISK-02 (current capacity), RISK-10 (ED vs RA) | §5 |
| PDMS optical window material | RISK-04 (CTE delamination) | §6 |
| Incoming inspection requirements | All above | §7 |

2.  LED EMITTERS — 660 nm AND 808–830 nm

⚠  RISK-08 (HIGH): Commodity LED emitters ship in Vf bins of ±0.20–0.30 V as standard. This width causes current imbalance between parallel strings even with per-string constant-current drivers, reducing dose accuracy and accelerating premature LED failure. The specifications below MUST appear on every purchase order.

2.1  Forward Voltage (Vf) Bin Specification

This is the most critical procurement specification in this document. Vf bin width directly determines current matching between LED strings. Wide bins cause current hogging even when per-string constant-current drivers are used, because the driver's compliance headroom is consumed unevenly across strings.

| Parameter | 660 nm LED | 808–830 nm LED |
| --- | --- | --- |
| Nominal forward voltage at 150 mA, 25°C | 2.0–2.2 V (confirm from datasheet) | 1.6–1.8 V (confirm from datasheet) |
| MAXIMUM Vf bin width — within a single purchase order | ± 0.10 V | ± 0.10 V |
| MAXIMUM Vf bin width — between purchase orders (lot-to-lot) | ± 0.15 V | ± 0.15 V |
| Vf measurement condition | 150 mA DC, junction at 25°C, 10 ms pulse to avoid self-heating | Same |
| Vf distribution data required with each lot | Yes — mean, σ, min, max across sampled units | Yes |
| Minimum sample size for Vf distribution certificate | 50 units per delivery lot | 50 units per delivery lot |

⚠  The ±0.10 V within-order bin width is tighter than commodity specification. Major LED suppliers (Seoul Semiconductor, Osram, Luminus) can meet this with explicit PO specification — but it must be stated. Without it, the supplier ships from mixed bin stock. Confirm bin availability and any price premium before first order.

2.1.1  Purchase Order Language — Vf Bin

The following text must appear verbatim in the "Special Requirements" or "Additional Specifications" section of every LED emitter purchase order:

"All units in this order shall be from a single forward voltage (Vf) bin. Maximum Vf spread within the order: ±0.10 V, measured at [150 mA DC / 10 ms pulse / junction at 25°C]. Supplier must provide a Vf distribution certificate (mean, standard deviation, minimum, maximum) based on a minimum sample of 50 units from this lot. Orders shipped without a Vf distribution certificate will be rejected at incoming inspection."

2.2  Wavelength Bin Specification

| Parameter | 660 nm LED | 808–830 nm LED |
| --- | --- | --- |
| Peak wavelength range | 655–670 nm | 805–835 nm |
| Maximum wavelength spread within order | ± 5 nm | ± 10 nm |
| Measurement condition | At 150 mA, 25°C junction temperature | Same |
| Wavelength certificate required | Yes — peak λ mean and range | Yes |

Wavelength tolerance is less critical than Vf for current matching, but matters for dose metering calibration. The reference photodiode calibration coefficient (irradiance-per-unit photocurrent) is wavelength-dependent. If lot-to-lot wavelength shifts exceed ±10 nm, the factory calibration coefficient must be recalculated for that lot.

2.3  Thermal and Package Specification

| Parameter | Requirement | Notes |
| --- | --- | --- |
| Package | SMD 2835 or equivalent ceramic substrate | Must be compatible with FPC reflow process (peak 245°C, 30 s) |
| Junction-to-case thermal resistance (θjc) | < 15 °C/W | Required for IEC 60601 surface temperature compliance model |
| Maximum junction temperature (Tj_max) | ≥ 125°C | Operating headroom above 62°C throttle threshold |
| Forward current — absolute maximum (If_max) | ≥ 350 mA | Must exceed 180 mA operating max with 2× margin |
| Pulse current rating | Confirm ≥ 180 mA at ≤ 25% duty cycle, ≤ 1 ms pulse width | Required before BOM lock — see CLAUDE.md §13.4 |
| Rated lifetime L70 | ≥ 80,000 hours at 150 mA nominal | L70 = time to 70% of initial luminous flux |
| Moisture sensitivity level (MSL) | MSL 3 or better | Must be stored and handled per IPC/JEDEC J-STD-033 |
| RoHS compliance | Required | IEC 63000 / EU Directive 2011/65/EU |

⚠  Pulse current rating at 120–180 mA for both wavelengths has not yet been confirmed from supplier datasheets. This must be verified before first LED order is placed. See CLAUDE.md §13.4 open issue.

2.4  Incoming Inspection — LED Emitters

Every delivery lot must be inspected before use in production. Incoming inspection failure results in lot quarantine and supplier notification. Do not use quarantined lots without written disposition from Hardware Engineering.

| Inspection item | Method | Sample size | Accept / Reject criterion |
| --- | --- | --- | --- |
| Vf distribution verification | Measure Vf at 150 mA, 10 ms pulse, 25°C on sample units. Compare to supplier certificate. | 10 units per lot (minimum) | Accept: all measured Vf within ±0.10 V of lot mean stated on certificate. Reject: any unit outside ±0.15 V of lot mean, or certificate not provided. |
| Peak wavelength spot-check | Measure λ_peak on spectrometer at 150 mA, 25°C. | 5 units per lot | Accept: λ_peak within stated wavelength range ± 5 nm. Reject: any unit outside range. |
| Package and marking | Visual inspection vs. datasheet. | 100% of reel labels; 5 physical units | Reject: wrong part number, damaged tape, missing/illegible marking. |
| MSL date code | Check reel label for exposure time and floor life. | 100% of reels | Reject: floor life expired or not stated. |
| Certificate of conformance (CoC) | Verify CoC accompanies shipment and references PO number. | Per shipment | Reject shipment if CoC absent. |

2.5  Out-of-Specification Lot Disposition

If incoming inspection reveals a lot that does not meet the Vf bin specification, the following disposition process applies:

Quarantine the lot immediately. Label reels "HOLD — PENDING DISPOSITION". Do not allow use in production.

Notify the supplier within 2 business days with the measured data and the specific PO specification that was not met. Request a replacement lot from a conforming bin.

Hardware Engineering Lead assesses whether the out-of-spec lot can be used with design adjustment. Options: (a) Use if Vf spread is ≤ ±0.15 V and all strings within a zone module are sorted from the same narrow sub-bin — requires manual sorting process; (b) Use at reduced current (lower RSET value) to reduce current-hogging risk — reduces irradiance proportionally; (c) Return to supplier for replacement.

Document disposition decision in the lot traveller. Any "use-as-is" decision requires written sign-off from Hardware Engineering Lead and Quality.

If a supplier repeatedly ships out-of-spec lots, qualify an alternative supplier from the approved backup list (§2.6) before placing next order.

2.6  Approved Supplier List — LED Emitters

| Supplier | Status | Wavelengths covered | Contact / Notes |
| --- | --- | --- | --- |
| Seoul Semiconductor | Primary | 660 nm + 808–830 nm | Confirm Vf bin capability before first order |
| Osram (ams-OSRAM) | Primary | 660 nm + 808–830 nm | Confirm Vf bin capability before first order |
| Luminus Devices | Backup | 660 nm + 808–830 nm | Qualify if primary suppliers cannot meet ±0.10 V bin requirement |

No supplier is yet fully qualified. "Primary" status means the supplier has been identified as likely capable. Full qualification requires: (1) sample lot with Vf certificate; (2) incoming inspection pass; (3) prototype build and functional test. Qualification must be completed before production orders are placed.

3.  LED DRIVER IC — INFINEON BCR421W

⚠  RISK-07 (MITIGATED): BCR421U (150 mA max) and BCR421W (350 mA max) are different variants with the same package (SOT-323) and similar appearance. Only BCR421W is approved. BCR421U must never be substituted, even if BCR421W is on allocation.

| Parameter | Required specification |
| --- | --- |
| Approved part number | Infineon BCR421W (350 mA rated) ONLY |
| Explicitly excluded part numbers | BCR421U (150 mA max — NEVER substitute). Any other BCR42x variant requires Hardware Engineering written approval. |
| Package | SOT-323 (SC-70), 3-pin |
| Current setting | Via external RSET resistor — value TBD pending regulatory opinion letter. Do NOT populate RSET until RSET value is released by Hardware Engineering. |
| Backup / alternate (approved) | Zetex ZXLD381 (350 mA, SOT-23-3) ¦ Diodes Inc. AL5809 (350 mA, SOT-23-3). Both require RSET recalculation from respective datasheets before use. |
| Approved supplier | Infineon Technologies — authorised distributor only (Mouser, DigiKey, Arrow, Avnet) |
| Counterfeit risk | HIGH — SOT-23/SOT-323 linear regulators are heavily counterfeited. Purchase from authorised distributor only. Request certificate of authenticity for orders > 500 units. |
| RoHS compliance | Required |

Purchase Order Language — BCR421W

"Part number: Infineon BCR421W. Package: SOT-323. SUBSTITUTION NOT PERMITTED — BCR421U and all other BCR42x variants are explicitly excluded. If BCR421W is on allocation or unavailable, contact Hardware Engineering before placing an order for any alternative part. Purchase from authorised distributor only. Certificate of conformance required."

4.  ZONE MODULE CONNECTOR — HIROSE FH34S / JAE FF03

✓  RISK-01 (MITIGATED): Molex SlimStack removed from specification. Hirose FH34S is the correct connector family for this application.

| Parameter | Required specification |
| --- | --- |
| Primary approved part number | Hirose FH34S-20S-0.5SH(50) — 20-position, 0.5 mm pitch, back-flip lever ZIF, SMT |
| Backup approved part number | JAE FF03 series — 0.5 mm pitch, high-cycle FPC connector |
| Explicitly excluded families | Molex SlimStack (board-to-board — incompatible with FPC). Molex 52610 / 52745 / 52892 series (rated 20–30 cycles — below requirement). ANY 0.35 mm pitch connector (current per pin too high for power rails at 0.9 A RMS). |
| Rated insertion cycles | ≥ 1,000 — MUST BE CONFIRMED FROM FULL HIROSE FH34S DATASHEET before BOM lock. Product marketing states ≥1,000 cycles; confirm in test conditions table. |
| Contact current rating | ≥ 0.5 A per contact |
| Pitch | 0.5 mm (NOT 0.35 mm) |
| Actuation | Back-flip lever ZIF — zero insertion force, tool-free extraction |
| Mated height | 1.0 mm maximum |
| Operating temperature | −40°C to +85°C minimum |
| Approved supplier | Hirose Electric — authorised distributor (DigiKey, Mouser, Arrow) |
| RoHS compliance | Required |

⚠  Before placing the first connector order: pull the complete FH34S datasheet from hirose.com and confirm that the "Endurance" or "Insertion/Withdrawal cycles" specification in the test conditions table states ≥ 1,000 cycles. If the full datasheet does not confirm this, contact Hirose applications engineering for written confirmation before placing the order.

Incoming Inspection — Connector

Verify part marking against approved part number FH34S-20S-0.5SH(50).

Verify position count (20) and pitch (0.5 mm) on 5 sample units using calibrated pin gauge.

Perform 5 insertion/extraction cycles on 3 sample units. Verify lever actuates correctly and FPC tail seats without resistance. Measure contact resistance at pins 1 and 20 before and after: must be < 50 mΩ per contact.

Certificate of conformance from Hirose or authorised distributor must accompany each shipment.

5.  FPC FABRICATION REQUIREMENTS

⚠  RISK-10 (MEDIUM): FPC fabricators default to electrodeposited (ED) copper. ED copper fails under dynamic flex. This requirement MUST appear on the fabrication drawing AND in the purchase order.

| Parameter | Required specification | Default if not stated (WRONG) |
| --- | --- | --- |
| Copper type | ROLLED ANNEALED (RA) copper — mandatory | Electrodeposited (ED) copper |
| Copper weight | 1 oz (35 µm) on all layers | May be substituted with 0.5 oz if not specified |
| Minimum power trace width | ≥ 2.0 mm for all VLED and GND power rails — see NP-HW-FPC-001 §10.1 | Trace width not enforced — layout engineer's discretion |
| Substrate | Polyimide (PI), 25 µm | May substitute polyester (PET) — not acceptable |
| Adhesive | Acrylic or adhesiveless laminate — NO epoxy | Epoxy adhesive (lower thermal performance) |
| Surface finish | ENIG on signal pads; hard gold ≥ 0.5 µm (cobalt-alloyed) on connector tail pads | ENIG only — insufficient wear resistance on connector tail |
| Dynamic bend radius | ≥ 25 mm at all bend zones — must be stated on drawing | No bend radius enforcement |
| IPC class | Class 2 minimum | Class 1 |
| Flammability | UL 94 V-0 | Not specified |
| Certificate of conformance | Required per lot | Not provided unless requested |

Purchase Order Language — FPC Fabrication

"COPPER TYPE: ROLLED ANNEALED (RA) — ELECTRODEPOSITED (ED) COPPER IS NOT ACCEPTABLE AND WILL RESULT IN REJECTION OF THE ENTIRE LOT. COPPER WEIGHT: 1 OZ (35 µm) ON ALL LAYERS. MINIMUM POWER TRACE WIDTH: 2.0 mm ON ALL VLED AND GND RAILS AS SHOWN ON DRAWING. IPC CLASS 2. UL 94 V-0. CERTIFICATE OF CONFORMANCE REQUIRED."

Incoming Inspection — FPC

Request material certification confirming RA copper from FPC fabricator. Reject lot if only ED copper certification is provided.

Measure copper trace width on power rails using calibrated optical comparator or cross-section: minimum 2.0 mm on VLED_660, GND_660, VLED_808, GND_808 buses. Sample: 3 units per lot.

Verify connector tail gold thickness via XRF: ≥ 0.5 µm hard gold. Sample: 2 units per lot.

Perform 20 insertion/extraction cycles on 2 sample units through the production ZIF connector. Verify connector tail shows no delamination, pad lifting, or gold wear through to base metal.

6.  PDMS OPTICAL WINDOW MATERIAL

⚠  RISK-04 (HIGH): PDMS–polyimide CTE mismatch (15:1) requires specific bonding process (sputtered SiO₂ interlayer). Standard PDMS from any supplier is acceptable for the base material; the bonding process is the critical variable. See NP-HW-FPC-001 §9 for full bonding process specification.

| Parameter | Required specification |
| --- | --- |
| Base material | Polydimethylsiloxane (PDMS) |
| Approved formulation | Dow Sylgard 184 (primary) ¦ equivalent 10:1 base:curing agent formulation (backup — confirm optical transmission) |
| Mixing ratio | 10:1 base:curing agent by weight — must be recorded on batch traveller |
| Cure conditions | 80°C for 2 hours minimum — or 120°C for 1 hour. Do not cure at room temperature (insufficient crosslink density for thermal cycling resistance). |
| Optical transmission at 660 nm (through spec thickness) | > 92% — measure before bonding; reject batches below threshold |
| Optical transmission at 808–830 nm | > 94% |
| Shore hardness | 40–50 A (confirm from formulation datasheet) |
| Batch traceability | Lot number of base and curing agent must be recorded in batch traveller for each production run |
| Storage | Unmixed components: store at 4–8°C, away from UV. Mixed/uncured PDMS: use within 4 hours of mixing or discard. |
| Bonding process requirement | Sputtered SiO₂ interlayer (75 nm) on PI surface required before O₂ plasma activation. Simple O₂ plasma bond alone is NOT acceptable for production — see NP-HW-FPC-001 §9.2. |

PDMS is a commodity material. Supplier qualification is less critical than for electronic components. The bonding process (§9.2 of spec) is the quality gate — peel test per batch is the acceptance criterion, not PDMS material source.

7.  INCOMING INSPECTION SUMMARY

The following table summarises all incoming inspection requirements across all component categories. Inspection records must be retained for a minimum of 7 years for medical device traceability.

| Component | Critical inspection item | Reject criterion | Sample size |
| --- | --- | --- | --- |
| 660 nm LED | Vf distribution certificate provided | Missing certificate → reject shipment | Per shipment |
| 660 nm LED | Vf spread within ±0.15 V of lot mean | Any unit outside → quarantine lot | 10 units / lot |
| 660 nm LED | Peak wavelength 655–670 nm | Any unit outside ±5 nm of range | 5 units / lot |
| 808–830 nm LED | Vf distribution certificate provided | Missing certificate → reject shipment | Per shipment |
| 808–830 nm LED | Vf spread within ±0.15 V of lot mean | Any unit outside → quarantine lot | 10 units / lot |
| 808–830 nm LED | Peak wavelength 805–835 nm | Any unit outside ±10 nm | 5 units / lot |
| BCR421W | Part number marking confirms W variant (not U) | Any BCR421U found → reject entire lot | 10 units / lot |
| BCR421W | Purchase from authorised distributor | Grey market source → reject | Per PO |
| FH34S connector | Part number and position count (20) confirmed | Wrong count or part → reject | 5 units / lot |
| FH34S connector | ZIF lever actuates and releases correctly | Stiff or non-actuating → quarantine | 3 units / lot |
| FH34S connector | Contact resistance < 50 mΩ after 5 cycles | Any pin > 50 mΩ → quarantine lot | 3 units / lot |
| FPC | RA copper material certificate | ED copper certificate or none → reject lot | Per lot |
| FPC | Power trace width ≥ 2.0 mm on VLED/GND rails | Any rail < 2.0 mm → reject lot | 3 units / lot |
| FPC | Hard gold on connector tail ≥ 0.5 µm (XRF) | Any pad < 0.4 µm → quarantine | 2 units / lot |
| PDMS | Optical transmission ≥ 92% at 660 nm | Any batch < 90% → reject | 1 sample / batch |

8.  REVISION HISTORY

| Rev | Date | Author | Description |
| --- | --- | --- | --- |
| A | 2026-05-06 | Claude / NeuroPulse Engineering | Initial release. Addresses RISK-01 (connector), RISK-04 (PDMS), RISK-07 (BCR421W), RISK-08 (LED Vf binning), RISK-10 (RA copper) from NP-HW-FPC-001 Risk Register Rev A. |