NeuroPulse

Zone Module FPC Assembly
First Article Inspection (FAI) Checklist

NP-FAI-ZM-001  Rev A  |  2026-05-06  |  Pre-Production Draft
References: NP-HW-FPC-001 Rev C · NP-PROC-FPC-001 Rev A · NP-DRV-SHELL-001 Rev A
Risk mitigations consolidated: RISK-01, 02, 04, 05, 06, 07, 08, 10, 11, 12

| Build lot / serial: |  | FAI date: |  |
| --- | --- | --- | --- |
| Headset serial number: |  | Inspector name: |  |
| FPC lot number: |  | Reviewed by: |  |
| Shell revision: |  | Final result: | PASS / FAIL |

SCOPE: This checklist covers the first article of each zone module FPC assembly. ALL items must pass before production release. Any FAIL triggers a disposition review — no waiver without Engineering and Quality sign-off. Items marked [BLOCKING] must pass before any other FAI items proceed.

SECTION 1 — PRE-INSPECTION DOCUMENT CHECK

Verify all required documents and BOM revisions are current before physical inspection.

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-D01 | FPC specification NP-HW-FPC-001 Rev C on file | Document check | Rev C or later confirmed |  |  |
| FAI-D02 | Procurement requirements NP-PROC-FPC-001 Rev A on file | Document check | Rev A or later confirmed |  |  |
| FAI-D03 | Shell tooling design review NP-DRV-SHELL-001 Rev A fully signed off (all 8 signatories) | Document check | All sign-off boxes in §8 completed |  |  |
| FAI-D04 | BOM confirms BCR421W (not BCR421U) — verify against approved BOM revision | BOM review | BCR421W in all LED driver positions |  |  |
| FAI-D05 | BOM confirms Hirose FH34S-20S-0.5SH(50) — no Molex SlimStack or 0.35 mm pitch connectors | BOM review | FH34S or JAE FF03 only in connector positions |  |  |
| FAI-D06 | FPC fab drawing specifies RA copper (IPC-4204/11 Type I) — check drawing fab notes | Drawing review | 'ROLLED ANNEALED' or 'RA per IPC-4204/11' visible as fab note 1 |  |  |
| FAI-D07 | RA copper material certificate accompanies FPC lot | Certificate check | Cert references lot number and IPC-4204/11 Type I |  |  |
| FAI-D08 | LED Vf bin distribution certificate accompanies LED lot (660 nm and 808–830 nm) | Certificate check | Cert present, lot mean ± range within ±0.10 V spec |  |  |

SECTION 2 — INCOMING COMPONENT INSPECTION

Verify purchased components before assembly. Failures here prevent assembly from proceeding.

2a. LED Emitters (660 nm and 808–830 nm)  [RISK-08 Vf binning · RISK-10 RA copper ref]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-L01 | 660 nm LED: measure Vf on 10-unit sample at 150 mA / 10 ms pulse / 25°C | Benchtop measurement | All 10 within ±0.10 V of lot mean on certificate |  |  |
| FAI-L02 | 808 nm LED: measure Vf on 10-unit sample at 150 mA / 10 ms pulse / 25°C | Benchtop measurement | All 10 within ±0.10 V of lot mean on certificate |  |  |
| FAI-L03 | 660 nm peak wavelength: spectrometer check on 5-unit sample | Spectrometer | λ_peak 655–670 nm, within ±5 nm of lot certificate |  |  |
| FAI-L04 | 808 nm peak wavelength: spectrometer check on 5-unit sample | Spectrometer | λ_peak 805–835 nm, within ±10 nm of lot certificate |  |  |
| FAI-L05 | Reel label part number and date code check | Visual | Part number matches BOM; floor life not expired (MSL 3) |  |  |

2b. BCR421W LED Driver IC  [RISK-07 variant confusion]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-C01 | [BLOCKING] Part marking: confirm 'W' suffix on every sampled IC | Visual / magnifier (10×) | Marking reads BCR421W. Any BCR421U → reject entire lot immediately |  |  |
| FAI-C02 | Package type: SOT-323 (SC-70) 3-pin | Visual / caliper | SOT-323 body size 2.0 × 1.25 mm confirmed |  |  |
| FAI-C03 | Authorised distributor CoC present | Certificate check | Distributor is Mouser, DigiKey, or Avnet; CoC on file |  |  |

2c. Hirose FH34S Connector  [RISK-01 connector family]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-K01 | [BLOCKING] Part number on reel label: confirm FH34S-20S-0.5SH(50) | Visual | Exact part number confirmed. Any Molex or 0.35 mm connector → reject |  |  |
| FAI-K02 | Position count: 20 contacts | Visual / pin count | Count = 20 |  |  |
| FAI-K03 | Lever actuation: back-flip lever opens and closes smoothly | Manual test, 3 samples | No stiffness, no binding, lever fully releases at 180° |  |  |
| FAI-K04 | Contact resistance: insert FPC tail, measure all 20 contacts | 4-wire measurement | All contacts < 50 mΩ after 5 insertion cycles |  |  |

2d. FPC Laminate  [RISK-10 RA copper · RISK-02 power trace]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-F01 | [BLOCKING] RA copper material certificate: confirm 'Rolled Annealed' per IPC-4204/11 Type I | Certificate check | Certificate explicitly states RA. 'ED' certificate → reject lot |  |  |
| FAI-F02 | Copper weight: 1 oz (35 µm ± 10%) confirmed on certificate | Certificate check | 35 µm ± 3.5 µm |  |  |
| FAI-F03 | Connector tail hard gold thickness: XRF measurement on 2 samples | XRF (2 pads/unit) | ≥ 0.5 µm gold. Any pad < 0.4 µm → quarantine lot |  |  |
| FAI-F04 | IPC class 2 certificate accompanies lot | Certificate check | Certificate references IPC-A-600 Class 2 or better |  |  |

SECTION 3 — ASSEMBLED ZONE MODULE INSPECTION

Inspect first-article assembled zone module PCB before installation into shell.

3a. Power Trace and Thermal Verification  [RISK-02 current capacity]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-T01 | Power trace width on VLED rail: measure from top layer photograph (calibrated image scale) | Optical / caliper | ≥ 2.0 mm on all VLED/GND power bus segments |  |  |
| FAI-T02 | Thermal imaging: run all 10 strings in one zone at full rated current (180 mA/string) for 60 s; capture IR image | FLIR or equivalent IR camera | No hotspot > 62°C junction temperature; uniform temperature gradient across strings |  |  |
| FAI-T03 | Trace continuity on all 20 power pins: measure resistance pin-to-string LED | 4-wire measurement | Per-trace resistance < 50 mΩ for any signal trace, < 10 mΩ for power rail |  |  |

3b. LED Driver Current Verification  [RISK-05 current hogging · RISK-07 BCR421W]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-E01 | [BLOCKING] Per-string current: measure actual current in each of 10 strings for one wavelength at rated RSET | Clamp meter / shunt, each string individually | All 10 strings within ±5% of target (e.g. 180 mA ± 9 mA). Any string outside ±10% → investigate driver |  |  |
| FAI-E02 | Across-wavelength match: compare 660 nm string mean vs 808 nm string mean current | As above, both wavelengths | Means within ±3% of each other at same RSET (same driver part) |  |  |
| FAI-E03 | BCR421W thermal rise: measure IC case temperature at rated current for 60 s (no heat sink) | NTC or contact thermocouple | Case temperature ≤ 85°C. Above 80°C → evaluate thermal derating |  |  |

3c. Reference Photodiode Optical Isolation  [RISK-06 photodiode contamination]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-P01 | [BLOCKING] Dark-field isolation test: cover all LEDs with opaque black tape; run photodiode readout in synchronous lock-in mode | Bench test, firmware lock-in mode | Photodiode ADC output < 1% of full-scale in absence of LED illumination |  |  |
| FAI-P02 | Response test: uncover LEDs; run all 10 strings in zone at 25% duty cycle; verify photodiode responds | Bench test | Photodiode output increases ≥ 90× compared to dark reading; signal stable across 10 s |  |  |
| FAI-P03 | LED aging simulation: attenuate LED output 30% with ND filter over PDMS window; verify photodiode detects reduction | ND filter + bench test | Reported J/cm² decreases by 28–32% (within ±2% of filter attenuation) |  |  |
| FAI-P04 | PDMS window optical transmission: measure before and after assembly with integrating sphere or photodiode | Integrating sphere / benchtop | ≥ 92% at 660 nm; ≥ 94% at 808–830 nm |  |  |

3d. PDMS Window Adhesion  [RISK-04 CTE delamination]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-M01 | Visual inspection of PDMS–PI bond line: no delamination, no air bubbles > 0.5 mm, no discolouration | Visual, 10× magnifier | No defects visible. Any delamination ≥ 1 mm → reject |  |  |
| FAI-M02 | Peel test (witness coupon from same batch): 90° peel test on 25 mm × 25 mm coupon | Instron or similar, 90° peel | Peel force ≥ 50 N/m. Below 20 N/m → reject batch (SiO₂ interlayer failure) |  |  |
| FAI-M03 | Thermal cycle (accelerated): 5 cycles −20°C to +70°C on witness coupon; re-inspect bond line | Thermal chamber | No new delamination after 5 cycles. If delamination: halt production, escalate to process engineer |  |  |

SECTION 4 — SHELL INSTALLATION AND ALIGNMENT VERIFICATION

RISK-12 primary mitigation. All 5 zone modules must be verified individually. The ±0.5 mm system alignment requirement is achieved via ±0.3 mm shell tooling tolerance (±0.2 mm margin).

4a. Zone Module Positional Accuracy  [RISK-12 alignment tolerance]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-A01 | [BLOCKING] Zone 1 (frontal left — Fp1/F3): CMM measurement of LED array centroid position vs. shell datum | CMM / optical profilometer | LED array centroid within ±0.3 mm of nominal Fp1/F3 target position in X and Y |  |  |
| FAI-A02 | [BLOCKING] Zone 2 (frontal right — Fp2/F4): CMM measurement | CMM | Within ±0.3 mm of nominal Fp2/F4 target |  |  |
| FAI-A03 | [BLOCKING] Zone 3 (central left — C3): CMM measurement | CMM | Within ±0.3 mm of nominal C3 target |  |  |
| FAI-A04 | [BLOCKING] Zone 4 (central right — C4): CMM measurement | CMM | Within ±0.3 mm of nominal C4 target |  |  |
| FAI-A05 | [BLOCKING] Zone 5 (parietal — P3/P4): CMM measurement | CMM | Within ±0.3 mm of nominal P3/P4 target |  |  |
| FAI-A06 | Zone-to-zone relative position: confirm inter-zone spacing matches 10-20 system inter-electrode distances | CMM / derived from A01–A05 measurements | All inter-zone distances within ±0.5 mm of 10-20 system nominal |  |  |
| FAI-A07 | Zone module stiffener datum features (reference holes or edges): verify present and within ±0.1 mm of nominal | CMM | Datum features present; position ±0.1 mm |  |  |
| FAI-A08 | Zone module re-insertion repeatability: remove and re-insert each zone module 3×; re-measure CMM position | CMM, 3 repeats per zone | Position repeatability ±0.1 mm (module self-locates to same position each insertion) |  |  |
| FAI-A09 | Zone module tactile identification: verify raised numeral (1–5) and braille zone abbreviation are present and readable on all 5 zone module bodies | Physical inspection; finger-reading test by 3 test subjects with eyes closed | Numeral present on all 5 modules; braille readable by each test subject (confirmed match to zone position). Any absent or unreadable → reject. |  |  |
| FAI-A10 | Shell tactile dot patterns: verify N raised dots adjacent to each zone slot for zones 1–5; confirm pattern is finger-findable from the slot edge | Physical inspection; blind-condition finger test by 3 test subjects | Correct dot count at each slot; dots reachable within 5 s by test subjects without visual reference |  |  |
| FAI-A11 | ZONE_ID electronic verification: read ADC voltage at pin 18 for all 5 zones and confirm correct resistor value per zone (ZM-01=10k, 02=22k, 03=47k, 04=100k, 05=220k ±5% ADC tolerance) | Hub MCU diagnostic mode (or bench ADC measurement) | All 5 zones read correct ID within ±5% band. Any zone outside band → reject module lot. |  |  |
| FAI-A12 | [BLOCKING] Wrong-zone ENABLE block: insert ZM-02 module into ZM-01 slot; confirm firmware asserts FAULT, blocks ENABLE, and issues audio alert via bone conduction | Live firmware test | ENABLE held low within 200 ms of module insertion. Bone conduction announces 'Wrong module' within 3 s. Session cannot start. |  |  |
| FAI-A13 | Correct-zone audio confirmation: insert all 5 modules correctly; confirm bone conduction announces correct zone name for each zone during headset initialisation sequence | Live firmware test; listen to bone conduction output | All 5 zone names announced correctly and in sequence. Audio intelligible at normal ambient noise level (≤60 dB background). |  |  |
| FAI-A14 | Mechanical key exclusion test: attempt to insert each zone module into each wrong slot (ZM-01 → slots 2–5, etc.); confirm mechanical key prevents insertion in all incorrect slot/module combinations | Physical manual test — 20 wrong-slot attempts per unit | Zero incorrect insertions possible. Module must be fully stopped by mechanical key before FPC tail contacts ZIF. Any mis-insertion possible → halt, escalate. |  |  |

4b. FPC Routing Path Verification (in-shell)  [RISK-11 bend radius]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-R01 | Zone 1 FPC: visual inspection of routing path in assembled headset — no visible kink, tight bend, or pinch point | Visual + endoscope (if needed) | No kink visible; FPC follows channel smoothly |  |  |
| FAI-R02 | Zone 2 FPC: same as FAI-R01 (highest risk — compound frontal curvature) | Visual + endoscope | No kink; smooth channel follow |  |  |
| FAI-R03 | Zones 3–5 FPC: visual inspection | Visual | No kink on any zone |  |  |
| FAI-R04 | FPC exit bend (Segment A) — physical radius check with radius gauge: all 5 zones | Radius gauge, bend check tool | ≥ 25 mm at all Segment A points. Any zone < 20 mm → halt and escalate |  |  |
| FAI-R05 | Zone 1 insertion/removal cycle test: insert/remove 10×; inspect FPC for crease, whitening, or deformation | Manual test + visual | No crease, whitening, or deformation after 10 cycles |  |  |
| FAI-R06 | Cross-zone interference: remove Zone 1 module; confirm Zone 2 FPC is not disturbed | Visual | Zone 2 FPC remains in channel during Zone 1 removal |  |  |
| FAI-R07 | ZIF lever access: confirm user can actuate ZIF lever (full 180° flip) without contacting FPC with fingertip | Manual test, 3 operators | All 3 operators actuate lever without FPC contact |  |  |

SECTION 5 — CONNECTOR LIFECYCLE VALIDATION

Zone modules are user field-replaceable. The Hirose FH34S connector is rated ≥1,000 cycles. FAI validates first-article connector performance at incoming (10-cycle) and production-release (50-cycle) levels.

5a. Connector Insertion Cycle Test  [RISK-01 connector family · RISK-11 routing path]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-K10 | Baseline contact resistance (before cycling): measure all 20 contacts | 4-wire, all 20 pins | All pins < 50 mΩ |  |  |
| FAI-K11 | Insert/remove zone module 50× (simulated field replacement pace — one cycle per 15 s) | Manual cycling, count logged | 50 cycles completed without mechanical failure of lever, body, or contacts |  |  |
| FAI-K12 | Post-cycle contact resistance: measure all 20 contacts | 4-wire, all 20 pins | All pins < 50 mΩ. Increase > 20 mΩ from baseline → flag for investigation |  |  |
| FAI-K13 | Post-cycle visual: inspect connector body and FPC tail for wear marks, gold wear-through, or fretting | Visual, 10× magnifier | No exposed base metal on contact pads; no cracking of FPC coverlay at tail |  |  |
| FAI-K14 | Post-cycle FPC position accuracy (re-run FAI-A01 to FAI-A05 on Zone 1 after 50 cycles) | CMM | Zone 1 position still within ±0.3 mm after 50 insertion cycles |  |  |

SECTION 6 — FUNCTIONAL SYSTEM TEST (ASSEMBLED HEADSET)

Run after all physical inspection items pass. Requires firmware-loaded headset connected to app.

6a. End-to-End Zone Function  [RISK-05 · RISK-06 · RISK-12]

| Item | Inspection / Test | Method | Accept Criterion | Result | Sign |
| --- | --- | --- | --- | --- | --- |
| FAI-SY01 | All 5 zones illuminate at 25% duty cycle — app confirms zone active, photodiode returns nonzero dose reading | App UI, firmware telemetry | All 5 zones report active; J/cm² > 0 for each zone |  |  |
| FAI-SY02 | Zone-selective deactivation: deactivate Zone 1 only; confirm Zones 2–5 remain active and Zone 1 extinguishes | App UI | Correct zone isolation confirmed |  |  |
| FAI-SY03 | PDMS fouling simulation: apply light smear of petroleum jelly to Zone 1 PDMS window; confirm photodiode detects reduced output and app flags maintenance reminder | App UI + petroleum jelly test | App shows reduced dose for Zone 1; maintenance reminder displayed within 2 sessions |  |  |
| FAI-SY04 | EEG signal quality on all 8 electrodes: confirm impedance < 10 kΩ with hydrogel tips installed | App impedance display | All 8 electrodes < 10 kΩ (green) |  |  |
| FAI-SY05 | Closed-loop session: run 5-minute EEG-adaptive PBM session; confirm adaptation events logged in UHDR | App session log | At least 1 adaptation event logged; UHDR session record created and encrypted |  |  |

SECTION 7 — NON-CONFORMANCE SUMMARY

List all items that did not pass during this FAI. All non-conformances must be dispositioned before production release — either by corrective action (re-inspect after fix) or by formal Engineering deviation approval. No BLOCKING item may be waived.

| FAI Item | Non-Conformance Description | Disposition | Corrective Action Ref | Closed by / Date |
| --- | --- | --- | --- | --- |

SECTION 8 — FAI RESULT AND SIGN-OFF

All checklist items in Sections 1–6 must show PASS, or non-conformances in Section 7 must be fully closed, before production release is authorised.

| Overall FAI result: | PASS / FAIL | Build lot approved for production: | YES / NO |
| --- | --- | --- | --- |
| Manufacturing Engineer: |  | Date: |  |
| Hardware EE Lead: |  | Date: |  |
| Quality / Reliability: |  | Date: |  |
| Programme Manager: |  | Date: |  |
| Notes: |  |  |  |

SECTION 9 — RISK MITIGATION CROSS-REFERENCE

All FAI items trace back to at least one risk register entry or specification requirement.

| Risk ID | Risk Title (abbreviated) | FAI Items | Status |
| --- | --- | --- | --- |
| RISK-01 | Connector family — Hirose FH34S | FAI-D05, FAI-K01–FAI-K04, FAI-K10–FAI-K14 | MITIGATED |
| RISK-02 | FPC power trace current capacity | FAI-T01, FAI-T02, FAI-T03 | MITIGATED |
| RISK-04 | PDMS–PI CTE mismatch delamination | FAI-M01, FAI-M02, FAI-M03, FAI-P04 | OPEN (process qualification pending) |
| RISK-05 | LED current hogging (per-string CCC) | FAI-E01, FAI-E02, FAI-E03, FAI-SY01 | MITIGATED |
| RISK-06 | Reference photodiode optical isolation | FAI-P01, FAI-P02, FAI-P03, FAI-SY03 | MITIGATED |
| RISK-07 | BCR421U vs BCR421W variant | FAI-C01, FAI-C02, FAI-D04 | MITIGATED |
| RISK-08 | LED Vf binning not on PO | FAI-L01–FAI-L05, FAI-D08 | MITIGATED |
| RISK-10 | ED copper default on FPC fab | FAI-D06, FAI-D07, FAI-F01, FAI-F02 | MITIGATED |
| RISK-11 | Shell bend radius not specified | FAI-R01–FAI-R07, FAI-K14 | MITIGATED |
| RISK-12 | Zone alignment tolerance not validated | FAI-A01–FAI-A08, FAI-SY04 | MITIGATED (this FAI) |
| RISK-15 | Zone keying — colour-blind + blind-accessible | FAI-A09, FAI-A10, FAI-A11, FAI-A12, FAI-A13, FAI-A14 | MITIGATED (this FAI) |

SECTION 10 — REVISION HISTORY

| Rev | Date | Author | Description |
| --- | --- | --- | --- |
| A | 2026-05-06 | NeuroPulse Engineering | Initial release. RISK-12 primary mitigation (alignment FAI). Consolidates FAI items from RISK-01, 02, 04, 05, 06, 07, 08, 10, 11, 12. 6 sections, 50+ inspection items. |