NeuroPulse Master Design Brief

Revision 5 — Closed-Loop Multi-Modal Neuromodulation Platform

Document Number: NP-DB-001 Rev 5

Date: May 16, 2026

Status: Pre-tooling design phase. No hardware committed yet.

Revision 5 — firmware architecture baselined; QMS established; 1064nm smart module complete;
T2 modalities firmware-specified; NPPS protocol language; hub control program;
iOS app roadmap; provisional accessories

# Section 1 — Key Additions Since Revision 4

The following table summarises all significant design decisions, firmware deliverables, hardware specifications, and QMS documents added or locked since Revision 4. Revision 4 covered the zone module engineering package through the 24-risk register. All items below are either LOCKED or explicitly marked PROVISIONAL or PENDING.

| # | Category | New decisions / documents added |
| --- | --- | --- |
| 1 | 1064nm Smart Zone Module (hardware complete) | Path B architecture locked. ATtiny402 I2C slave + 3× IRLML6344 N-FETs. 550 LEDs (200×660nm + 200×808nm + 150×1064nm). InGaAs PDs (Hamamatsu G12180-010A). EPITEX L1064-02AU emitters via Marubeni. BOM delta +$23–28/module. Retail $149–199/zone, $599–699 all-5-zone kit. NP-HW-FPC-001 Rev E, NP-FPC-ZM-SM-01, NP-TOOL-ZM-SM-001 Rev A. |
| 2 | Hub PCB Rev B | Vishay DG2788A dual-SPDT per-slot analog switch switches TIA Rf 47kΩ→22kΩ when smart module detected. NXP PCA9546A 4-ch I2C isolation switch. GAIN_SEL GPIO sequenced before I2C mux. Hub BOM delta $1.41–1.91. NP-HW-HUB-001 Rev B. |
| 3 | eMMC partition architecture + encryption | 9-partition GPT layout (8,192 MiB total). UHDR: AES-256-XTS, Argon2id key derivation (biometric/PIN, key never held by NeuroPulse). SHDR: AES-256-XTS, HKDF manufacturing root key. LittleFS on UHDR/SHDR/Config. 27-element session data classification table. NP-FW-EMMC-001 Rev A. |
| 4 | Dual-bank OTA bootloader firmware (written) | NXP i.MX RT1062 bootloader. Ed25519 signature verify (self-contained). 9-step OTA state machine. USB-C DFU 1.1 recovery. SNVS_LPGPR0 bank flag. firmware/bootloader/. G1-16 CLOSED. |
| 5 | HRV biofeedback firmware (written) | Six-module static library: ppg, coherence (256-pt Welch PSD), pacer, tavns_sync, eeg_biofeedback, session. Four protocols: standalone, HRV+taVNS synchronised, HRV+EEG dual, HRV+PBM. firmware/hrv_biofeedback/. G2-12 CLOSED. |
| 6 | Zone bone conduction announcement firmware (written) | RISK-15 Layer 5. LPADC1 zone detection with 3×100ms debounce. DDS audio from 256-entry Q15 sine LUT via SAI3 I2S. Five parallel slot state machines. firmware/zone_announce/. G2-10 CLOSED. |
| 7 | sLORETA-guided HD-tDCS firmware spec | Scalar sLORETA weight matrix W (2,447 voxels × 21 channels, ≈205KB LPSDR4). 4×1 ring montage. 30s ramp. 7 clinical targets. NP-FW-HD-001 Rev A. firmware/sloreta_hdtdcs/. G3-07 SOFTWARE BASELINED. |
| 8 | Cervical VNS firmware + 510(k) Q-Sub | Dual-processor cardiac interlock. HR change >15 BPM in 5s → GPIO cutoff <100ms. 3-module firmware. 510(k) predicate: electroCore gammaCore K163334+K173323. NP-FW-CVNS-001 Rev A, NP-REG-CVNS-001 Rev A. G3-08 SOFTWARE BASELINED. RISK-25 MITIGATED. |
| 9 | QMS established at company formation | Six documents released simultaneously: NP-QMS-001 (manual), NP-DHF-001 (DHF index), NP-QMS-DC-001 (design controls), NP-RM-001 (risk management plan), NP-SW-001 (IEC 62304 software development plan), NP-QMS-CAPA-001 (CAPA procedure). Software classes locked: Safety MCU = Class C; Main processor = Class B; App = Class B. Date: 2026-05-13. |
| 10 | Shell tooling spec | NP-TOOL-SHELL-001 Rev A. Four mandatory features: zone slot plug anchor posts (×5, colour-coded), accessory port cover anchor posts (×3), EEG cable routing channel (8×5mm outer CFRP), temporal wing anchor boss (PROVISIONAL — awaits HOPE Phase 3). |
| 11 | Lens tooling Rev B | NP-TOOL-LENS-001 Rev B. F-07 added: goggle arm anchor hook for lens rim guard tether, 4mm ID loop, superior-posterior corner, shade clearance ≥4mm (MR-19), tether ≤45mm posterior routing no FOV crossing (MR-20). |
| 12 | EEG cable routing CLOSED | OI-09 CLOSED. 8×5mm main trunk + 6×4mm lateral branches, outer CFRP inner surface. Hub → occipital arch (~120mm) → crown junction → left/right branches. 18–22mm FPC-to-EEG separation. DRC-18/22/23 CLOSED. G2-11 CLOSED. NP-DRV-SHELL-001 Rev B. |
| 13 | FPC layout frozen | NP-HW-FPC-001 Rev D (base) + Rev E (smart module). PD2 position locked: X=33.0mm, Y=39.0mm. 1.6mm annular ring, scalp-facing hard gold ≥0.5µm. OI-3 CLOSED. G1-15 CLOSED. |
| 14 | T2 combined session protocol | NP-SES-1064-001 Rev A. 1064nm zone modules + 1170nm laser coordinated. Three-tier penetration: 660nm surface / 1064nm cortical / 1170nm deep subcortical. Thermal throttle priority: 1170nm first → CH_C → CH_B. |
| 15 | RISK-03 scope expansion | NP-REG-PBM1064-001 Rev A. 12 questions to regulatory counsel covering 1064nm pathway, aggregate irradiance (600 mW/cm²), T2 combined session, depth-tier FTC claims. Marketing gate on all 1064nm irradiance claims until opinion letter received. |
| 16 | 1064nm bibliography addendum | NP-BIB-1064-001 Rev A. Six entries including Yao et al. Science Advances 2022 (n=20 RCT, right PFC 1064nm, working memory CDA improvement, primary evidence). |
| 17 | Hub control program (written) | NP-FW-HUB-001 Rev A. Four FreeRTOS tasks: module probe, protocol runner, telemetry, safety SPI heartbeat (200ms). Probes all 11 slots (zones 0–4, EEG, audio, visual, VNS+HRV, intranasal, cervical VNS). Ed25519 protocol verification. Routes telemetry to UHDR/SHDR per 27-element classification table. firmware/hub_control/. |
| 18 | NPPS protocol scripting language | NeuroPulse Protocol Script (.npps). UTF-8 text format for defining, sharing, and storing session protocols. 15 modality blocks (T1+T2+accessories). Interval, composite, and limits blocks. Parser + predefined protocol library (19 protocols). App-side compiler (hubCompiler.ts). docs/npps-reference.md. |
| 19 | T2 + accessory wire format | Hub wire format expanded: 6 new module type IDs (QEEG_21CH, TMS, PBM_1170NM, CLIN_TACS, HD_TDCS, VIBROTACTILE). mods_active_mask uint8_t → uint32_t. Stub detect/init registered for all T2 modules. firmware/hub_control/. |
| 20 | iOS app roadmap | NP-APP-ROADMAP-001 Rev A. Apple Watch sync app as Year 1 post-launch deliverable. 4 phases: session status + HRV breathing ring (Month 1–2), audio sync (Month 3–4), haptic sync 40Hz (Month 4–5), visual flicker (Month 6+). docs/neuropulse_ios_app_roadmap_001.md. |
| 21 | Provisional accessories | 40Hz vibrotactile mastoid LRA pad (PROVISIONAL — await HOPE Phase 3 mid-2026, Cognito Therapeutics n=670). Temporal wing anchor boss provisioned in NP-TOOL-SHELL-001 Rev A F-04 at zero incremental tooling cost. Apple Watch sync app (3 sync channels, PROVISIONAL). BOM delta $0 for Watch app. |
| 22 | Risk register: 25 risks | RISK-25 (cervical VNS cardiac reflex) added and MITIGATED via dual-processor firmware interlock. Total: 25 risks (RISK-01 through RISK-25); 23 MITIGATED; RISK-03 and RISK-20 remain OPEN (external). |

# Section 2 — 1064nm Smart Zone Module (Complete Specification)

The 1064nm smart zone module is the most significant hardware addition since Revision 4. It extends the base PBM zone module to a tri-wavelength device using an on-module driver IC architecture (Path B). The module is field-installable by the user via the existing 20-pin FPC connector and zone slot, and is mechanically keyed to prevent insertion into hubs without the required Rev B TIA gain switch circuitry.

## 2.1 Architecture — Path B (On-Module Driver)

Documents: NP-HW-FPC-001 Rev E, NP-FPC-ZM-SM-01, NP-TOOL-ZM-SM-001 Rev A, NP-FW-PBM1064-001 Rev A

| Parameter | Specification |
| --- | --- |
| On-module driver IC | Microchip ATtiny402 (I2C slave, SOT-23-8, address 0x30, UPDI-programmed before assembly) |
| LED drivers | 3× Infineon IRLML6344TRPBF N-channel MOSFETs (SOT-23, RDS(on)=27mΩ, fully enhanced at 3.3V gate drive) |
| Channel allocation | CH_A = 660nm ¦ CH_B = 808nm ¦ CH_C = 1064nm |
| LED count | 550 total: 200×660nm + 200×808nm + 150×1064nm (tri-wavelength interspersed pattern) |
| Module footprint | 66×78mm — unchanged from base module (shared chassis, single production mould variant) |
| 1064nm emitters | EPITEX L1064-02AU via Marubeni |
| InGaAs photodiodes | Hamamatsu G12180-010A (0.90 A/W at 1064nm, <1nA dark current) — PD1 (forward emission) + PD2 (scalp-side backscatter) |
| PD position | Annular ring, both PD1 and PD2 at X=33.0mm, Y=39.0mm from module reference corner (geometric centre of LED array) |
| PD contact (PD2) | 1.6mm annular ring, scalp-facing copper layer, hard gold ≥0.5µm — pin 19 |
| Rigidizer sub-board | 22×14mm, carries ATtiny402 + 3 FETs |
| Driver BOM | $1.00 total (ATtiny402 $0.38 + 3× IRLML6344 $0.36 + passives) |
| Module BOM delta vs base | +$23–28 per module |
| Retail pricing | $149–199 per zone module / $599–699 all-5-zone kit |
| FPC document | NP-FPC-ZM-SM-01 (smart module layout variant of NP-HW-FPC-001 Rev E) |

## 2.2 Smart Module Detection and Hub Interface

| Item | Detail |
| --- | --- |
| ZONE_ID resistor | 3.3kΩ (1%) on pin 18 → ADC < 1100 counts (12-bit, 10kΩ pull-up, 3.3V ref) |
| Base module ZONE_ID range | 10kΩ–220kΩ → ADC 1100–4000 counts (ZM-01 through ZM-05) |
| Hub I2C bus | LPI2C3, 400kHz fast mode — 5 independent per-slot buses |
| ATtiny402 I2C address | 0x30 — isolated per-slot by NXP PCA9546A switch (address 0x70) |
| ACK timeout | 5ms after I2C bus enable — FAULT logged to SHDR if no ACK |
| ATtiny402 register map | 0x00 STATUS ¦ 0x01 CH_ENABLE ¦ 0x02/03/04 CUR ¦ 0x05/06/07 PWM_FREQ ¦ 0x08/09/0A DUTY ¦ 0x0B THERMAL ¦ 0x0C FAULT_LATCH ¦ 0x0D CONFIG |
| Detection sequence | ZONE_ID debounce (3×100ms, ≥2/3 pass) → GAIN_SEL assert → 10µs settle → I2C mux enable → I2C probe |

## 2.3 Hub PCB Rev B Requirement (OI-PBM-HW-01 — SPECIFIED)

The InGaAs PD has 0.90 A/W responsivity at 1064nm. At 72µA photocurrent, the existing 47kΩ TIA feedback resistor produces 0.90 × 72µA × 47kΩ = 3.38V — exceeding the 3.3V ADC rail and saturating the TIA. Hub PCB Rev B is required before any optical bench FAI.

| Component | Function | Qty | BOM |
| --- | --- | --- | --- |
| Vishay DG2788A dual-SPDT analog switch | Switch TIA Rf 47kΩ→22kΩ when smart module detected (GAIN_SEL GPIO HIGH after ZONE_ID debounce) | 5 (one per slot) | $0.75–1.25 total |
| NXP PCA9546A 4-ch I2C switch | Isolate slots 0–3 sharing ATtiny402 address 0x30 (slot 4 via GPIO mux) | 1 | $0.41–0.51 |
| 4.7kΩ pull-ups | I2C bus pull-ups for 5 independent LPI2C3 buses | 12× | ~$0.10 |
| 47kΩ / 22.1kΩ resistors | TIA feedback network per slot (10 each) | 20× | ~$0.15 |

Hub BOM delta total: $1.41–1.91. Hub PCB Rev B Gerber build (OI-HUB-03) is BLOCKING for FAI-SM-04 and FAI-SM-06 optical bench tests.

## 2.4 Smart Module Tooling Variant (NP-TOOL-ZM-SM-001 Rev A)

A mould variant is required for the smart module's rigidizer sub-board and asymmetric mechanical key. All other zone module features (keying, gasket, eject lever, PD2 aperture) are inherited unchanged.

| Moulded Feature | Specification | Purpose |
| --- | --- | --- |
| Rigidizer cavity | 24×16×3.5mm internal recess | Houses ATtiny402 + FET sub-board without increasing total module height |
| Asymmetric mechanical key | +2.0mm lateral face width vs base module | Prevents physical insertion into hub slots lacking Rev B TIA gain switch — prevents TIA saturation if smart module inserted into unmodified hub |
| All other features | Identical to NP-TOOL-ZM-001 Rev A (F-01 through F-08) | Shared mould tooling approach — only cavity and key differ |

## 2.5 Safety and Dose Limits

| Parameter | Limit | Enforcement |
| --- | --- | --- |
| Duty cycle ceiling | 25% maximum (DUTY_MAX_REG = 0x32) | Hub firmware layer — unconditional, cannot be overridden by app or session protocol |
| Session dose — 660nm / 808nm | 60 J/cm² per zone per channel per session | np_pbm1064 session monitor — abort on breach |
| Session dose — 1064nm | 36 J/cm² per zone per session | np_pbm1064 session monitor — abort on breach |
| Aggregate irradiance ceiling | 600 mW/cm² three-channel combined (OI-PBM-05) | PENDING RISK-03 counsel Q5/Q6 — marketing gate maintained |
| PD1/PD2 ratio disambiguation | Ratio narrows fouling vs LED aging failure mode | PD1↓ PD2 stable → fouling prompt. Both↓ proportional → LED aging correction (RISK-14 Option B) |
| Factory calibration | K_PD1_1064, K_PD2_1064, K_ratio_nom per zone per wavelength stored in Config partition | OI-PBM-04 OPEN — integrating sphere procedure required at manufacture |

## 2.6 T2 Combined 1064+1170nm Session (NP-SES-1064-001 Rev A)

In T2 configurations, 1064nm smart zone modules (cortical depth) are coordinated with the existing 1170nm laser system (subcortical depth) by the np_pbm1064_t2_combined session orchestrator. This creates a three-tier photobiomodulation penetration stack.

| Tier | Wavelength | Target depth | Source |
| --- | --- | --- | --- |
| Surface | 660nm | Scalp / superficial cortex | Zone module CH_A (200 LEDs per zone) |
| Cortical | 1064nm | ~20–30mm (cortical grey matter) | Smart zone module CH_C (150 LEDs per zone) |
| Deep subcortical | 1170nm | 35–40mm (subcortical structures) | T2 1170nm laser diodes with TEC stabilisation |

Thermal throttle priority (highest risk first): 1170nm laser throttled first, then CH_C (1064nm), then CH_B (808nm). CH_A (660nm) and all remaining channels throttled last. sLORETA coordination for T2 depression protocol requires T2 prototype (OI-SES-04 OPEN).

## 2.7 FAI Status

| FAI Item | Type | Status | Blocker |
| --- | --- | --- | --- |
| FAI-SM-01 | I2C communication protocol | Software PASS (PR #57) | None |
| FAI-SM-02 | ZONE_ID detection and debounce | Software PASS (PR #57) | None |
| FAI-SM-03 | Duty cycle ceiling enforcement | Software PASS (PR #57) | None |
| FAI-SM-04 | Optical power vs register (1064nm) | Hardware PENDING | Hub PCB Rev B Gerber build (OI-HUB-03) |
| FAI-SM-05 | PD1/PD2 ratio disambiguation logic | Software PASS (PR #57) | None |
| FAI-SM-06 | Dose metering accuracy (1064nm) | Hardware PENDING | Hub PCB Rev B Gerber build (OI-HUB-03) + calibrated optical bench |
| FAI-SM-07 | Thermal throttle priority | Hardware PENDING | T1 prototype |
| FAI-SM-08 | Smart module mechanical key insertion test | Hardware PENDING | Physical samples |
| FAI-SM-09 | Session dose limit abort | Software PASS (PR #57) | None |
| FAI-SM-10 | SHDR auth log on insert/remove | Software PASS (PR #57) | None |
| FAI-SM-11 | T2 combined session orchestration logic | Software PASS (PR #57) | None |

# Section 3 — Firmware Architecture Baselined

## 3.1 eMMC Partition Architecture (NP-FW-EMMC-001 Rev A)

The eMMC partition layout is baselined in NP-FW-EMMC-001 Rev A. All firmware storage architecture must conform to this specification — no storage code may be authored before this specification is reviewed (NP-COORD-001 G1-16 CLOSED). The 9-partition GPT user area totals 8,192 MiB across dual hardware boot partitions (bootloader) plus user area.

| Partition | Size | Filesystem | Encryption | Key holder | Notes |
| --- | --- | --- | --- | --- | --- |
| Firmware Bank A | 128 MiB | raw | None | — | Write-protected in production. Active running firmware. |
| Firmware Bank B | 128 MiB | raw | None | — | OTA staging target. Bootloader verifies Ed25519 before promoting. |
| Safety MCU FW Staging | 4 MiB | raw | None | — | Manual update only — never automated OTA. User confirmation + SPI readback required. |
| Config / Calibration | 16 MiB | LittleFS | AES-256-XTS (mfg key) | NeuroPulse (SNVS OTPMK) | Factory cal coefficients, ZONE_ID thresholds, TRNG salt for UHDR key derivation, OTA state record. |
| SHDR | 512 MiB | LittleFS | AES-256-XTS (mfg key) | NeuroPulse | Fleet telemetry. NeuroPulse decrypts for warranty / predictive maintenance. No user biology. |
| UHDR | 6,903 MiB | LittleFS | AES-256-XTS (Argon2id key) | User ONLY | ~84% of user area. NeuroPulse never holds or can reconstruct key. Biometric/PIN derived. |
| Scratch | 500 MiB | None (raw) | None | — | Zeroed on every boot. SANITIZE command issued. Used for OTA staging and research anonymisation workspace. No persistent data. |

Key cryptographic design decisions:

UHDR key derivation: Argon2id (64MB memory, 4 iterations, parallelism p=1) from biometric/PIN input → 64-byte output split into K1+K2 for AES-256-XTS. 32-byte TRNG salt stored in Config partition. NeuroPulse never holds, transmits, or can reconstruct the key.

SHDR key: HKDF-SHA256 from NeuroPulse manufacturing root key + eMMC CID hash → stored in Config partition under SNVS OTPMK fuse. NeuroPulse can decrypt for fleet analytics and warranty support.

Write endurance: EXT_CSD wear indicator + LittleFS block utilisation logged to SHDR every boot and every 100 sessions. PRE_EOL_INFO=0x03 blocks session start.

Dual-bank OTA: Ed25519 signature verified on all firmware images before promotion. 9-step OTA sequence with streaming SHA-256 over 128 MiB banks, write-with-readback verification, automatic rollback after 3 failed boot attempts of OTA candidate.

UHDR backup: nightly incremental backup on USB-C power, same Argon2id-derived key, to USB-C local or E2E encrypted cloud. NeuroPulse cannot decrypt backup archives.

## 3.2 Dual-Bank OTA Bootloader (firmware/bootloader/)

The first firmware code authored for NeuroPulse. Implements boot sequence B-1 through B-9 per NP-FW-EMMC-001 Rev A §8.1–§8.4. Targets NXP i.MX RT1062 (Cortex-M7). Linker script targets OCRAM at 0x20200000 with 48 KiB bootloader limit. CMake build system with arm-none-eabi toolchain.

| Module | Function |
| --- | --- |
| np_main.c | IVT, vector table relocation, hardware initialisation, top-level boot flow B-1 through B-9 |
| np_boot_selector.c | SNVS_LPGPR0 bank flag management: bits [1:0] BOOT_BANK, [7:2] BOOT_ATTEMPTS (max 3), [8] OTA_PENDING, [16] DFU_FORCED |
| np_emmc.c | USDHC2 eMMC driver: partition switching (hardware boot parts 1/2/user area), 8-bit bus, HS timing, write-with-readback |
| np_signature.c | Self-contained Ed25519 verify (RFC 8032 §5.1.7) + SHA-256 + SHA-512 + CRC32 + GF(2²⁵⁵−19) field arithmetic. Constant-time final comparison. No external crypto library dependency. |
| np_ota.c | 9-step OTA state machine. Streaming sector-by-sector SHA-256 (avoids loading full 128 MiB into SRAM). np_ota_state_t persisted to Config partition at two mirrored sectors. Rollback on 3 failed boot attempts. |
| np_dfu.c | USB DFU 1.1 class on USB1 (USB-HS OTG). DFU_DNLOAD/MANIFEST with Ed25519 verification before eMMC write. Writes to Scratch, verifies, then copies to inactive bank. No unauthenticated firmware can be flashed. |

Manufacturing root public key is a placeholder in np_signature.c — MUST be replaced at secure build step before any production image is signed.

## 3.3 Hub Control Program (NP-FW-HUB-001 Rev A, firmware/hub_control/)

Main application firmware for the NXP i.MX RT1062 hub processor running FreeRTOS. Four top-level tasks manage all hub functions from powerup through session completion.

| Task | Function |
| --- | --- |
| Module probe task | On powerup, scans all 11 physical slots and accessory ports: zone 0–4, EEG, audio, visual, VNS+HRV, intranasal, cervical VNS. Populates np_module_registry — a flat array of init/control/telemetry/shutdown function pointers. T2 module stubs (QEEG_21CH, TMS, PBM_1170NM, CLIN_TACS, HD_TDCS, VIBROTACTILE) return NOT_PRESENT until hardware drivers are built. |
| Protocol runner task (np_session_runner) | Receives Ed25519-signed compiled binary protocol from app (or runs stored autonomous protocol). Insertion-sort commands by start_ms. Dispatches at session-relative times. Issues implicit stop at (start_ms + duration_ms). Device-serial replay guard in session header prevents cross-device protocol reuse. |
| Telemetry task (np_session_log) | Collects per-module telemetry every second. Routes all telemetry to UHDR or SHDR per the 27-element classification table (NP-FW-EMMC-001 Rev A §12). EEG sample blocks written directly to UHDR partition via HAL to avoid buffering 12 KB/s in SRAM. |
| Safety SPI heartbeat task (np_safety_spi) | 200ms SPI heartbeat to STM32G071 safety MCU carrying requested-enable bitmask. Safety MCU owns all stimulation GPIO lines and cuts them on 1,500ms watchdog expiry if heartbeat stops. Implements IEC 62304 Class C interlock. |

mods_active_mask expanded from uint8_t to uint32_t to accommodate NP_MOD_VIBROTACTILE at bit 16 (>8 bits required).

## 3.4 NPPS Protocol Scripting Language (docs/npps-reference.md)

NeuroPulse Protocol Script (.npps) is a UTF-8 text format for defining, sharing, and storing session protocols. It is the canonical format for the predefined protocol library and for user-created custom protocols.

| Feature | Detail |
| --- | --- |
| File format | UTF-8 text, extension .npps, human-readable and diff-friendly |
| Modality blocks (15) | T1: pbm_transcranial, pbm_intranasal, eeg, bes, tdcs, vns_hrv, audio, visual T2: qeeg_21ch, tms, pbm_1170nm, clinical_tacs, hd_tdcs, cervical_vns Accessory: vibrotactile_40hz |
| Interval blocks | Phase sequence within a session (ramp_up, steady, ramp_down) with per-phase parameter overrides |
| Composite block | Combines multiple .npps protocols with timing offsets into a coordinated multi-modal session |
| Limits block | Per-modality dose constraints referenced by protocol validator before session start |
| Predefined library | 19 protocols covering all T1 configurations and key T2/composite use cases |
| App-side compiler | hubCompiler.ts: converts NPProtocolDefinition → 64-byte header + binary command records + Ed25519 signature placeholder |
| Device-side validation | Hub verifies Ed25519 signature before executing any protocol — rejects unsigned or corrupted sessions |

## 3.5 All Firmware Modules Written to Date

| Module | Directory | Specification | Gate Status |
| --- | --- | --- | --- |
| Dual-bank OTA bootloader | firmware/bootloader/ | NP-FW-EMMC-001 Rev A §8 | G1-16 CLOSED |
| HRV biofeedback protocol | firmware/hrv_biofeedback/ | NP-FW-HRV-001 Rev A | G2-12 CLOSED |
| Zone bone conduction announcement | firmware/zone_announce/ | NP-FW-ZA-001 Rev A | G2-10 CLOSED |
| sLORETA-guided HD-tDCS | firmware/sloreta_hdtdcs/ | NP-FW-HD-001 Rev A | G3-07 SOFTWARE BASELINED; hardware FAI PENDING |
| Cervical VNS safety interlock | firmware/cervical_vns/ | NP-FW-CVNS-001 Rev A | G3-08 SOFTWARE BASELINED; hardware FAI PENDING |
| 1064nm smart zone module | firmware/pbm_1064nm/ | NP-FW-PBM1064-001 Rev A | FAI-SM-01/02/03/05/09/10/11 PASS; SM-04/06/07/08 hardware PENDING |
| Hub control program | firmware/hub_control/ | NP-FW-HUB-001 Rev A | T1 modules complete; T2 stubs present — drivers PENDING hardware |

# Section 4 — Quality Management System (QMS)

QMS established at company formation: 2026-05-13. Six QMS documents released simultaneously. Scope: T1 + T2, all modalities, all firmware, iOS/Android app. Regulatory framework: 21 CFR Part 820 / ISO 13485:2016 / ISO 14971:2019 / IEC 62304 / IEC 62366-1.

All pre-formation design documents (RISK-01..25, all NP-* specifications, CLAUDE.md Rev 1–8) retroactively entered under QMS change control at establishment. DHF baseline established.

## 4.1 IEC 62304 Software Safety Classifications (locked)

| Software Item | Processor | Classification | Rationale |
| --- | --- | --- | --- |
| SW-01 Safety MCU firmware | STM32G071, bare-metal | CLASS C | Owns all stimulation GPIO enable lines. No independent safety backstop exists above it. Single-point failure can cause patient harm. |
| SW-02 Main processor firmware | i.MX RT1062, FreeRTOS | CLASS B | Safety MCU provides independent hardware backstop with SPI watchdog. App crash cannot cause unsafe stimulation. |
| SW-03 iOS/Android app | iPhone/Android, managed runtime | CLASS B | Safety MCU provides independent hardware backstop. App crash cannot cause unsafe stimulation. |

MISRA C:2012 mandatory for SW-01 (Class C). Advisory for SW-02. V-model development for SW-01; iterative V-model for SW-02/SW-03.

## 4.2 QMS Documents Released

| Document | Number | Status | Contents |
| --- | --- | --- | --- |
| QMS Manual | NP-QMS-001 Rev A | RELEASED 2026-05-13 | Quality policy, process map, management responsibility (Quality Lead position open; interim CEO), eQMS formalisation roadmap (Month 3–24). Establishes NeuroPulse QMS per 21 CFR Part 820 + ISO 13485:2016. |
| Design History File Index | NP-DHF-001 Rev A | RELEASED 2026-05-13 | Master DHF index per 21 CFR §820.30(j). All pre-formation documents retroactively entered. DHF completeness assessment. Firmware source code directories declared as design output records. Planned additions: NP-DP-001, NP-DT-001, NP-HFE-001, NP-PMS-001. |
| Design Controls Procedure | NP-QMS-DC-001 Rev A | RELEASED 2026-05-13 | 21 CFR §820.30 compliance. Covers design planning, inputs, outputs, review, verification, validation, transfer, and change control. Significant change order process vs minor change notification defined. |
| Risk Management Plan | NP-RM-001 Rev A | RELEASED 2026-05-13 | ISO 14971:2019. Severity S1–S5 / probability P1–P5. 5×5 risk acceptability matrix (ACCEPTABLE / ALARP / UNACCEPTABLE). RISK-01..25 under formal change control. Risk control hierarchy: inherent safe design → protective measures → information for safety. |
| Software Development Plan | NP-SW-001 Rev A | RELEASED 2026-05-13 | IEC 62304:2006+AMD1:2015. Safety response time requirements: heartbeat→cutoff ≤50ms; cardiac interlock ≤100ms; photoparoxysmal ≤200ms. SOUP tracking (FreeRTOS, LittleFS, Ed25519, CMSIS-DSP). OTA update: SW-01 never automated. |
| CAPA Procedure | NP-QMS-CAPA-001 Rev A | RELEASED 2026-05-13 | 21 CFR §820.100. 8 trigger sources. Critical/Major/Minor classification with response SLAs. 8-step CAPA process. MDR reporting obligations (30-day for death/serious injury, 5-day imminent hazard). Record retention 5 years minimum. |

## 4.3 Planned QMS Deliverables

| Document | Number | Target Month | Requirement |
| --- | --- | --- | --- |
| Design and Development Plan | NP-DP-001 | Month 3 | Formal per 21 CFR §820.30(b). Currently CLAUDE.md + NP-COORD-001 serve as interim plan. |
| Design Input/Output Traceability Matrix | NP-DT-001 | Month 6 | Links all requirements to verification evidence. Required for DHF completeness. |
| Human Factors Engineering Plan | NP-HFE-001 | Month 9 | IEC 62366-1 / FDA HFE Guidance 2016. URRA + formative + summative testing plan. Required for 510(k). |
| Post-Market Surveillance Plan | NP-PMS-001 | Month 12 | Required before T1 launch and for ongoing EU MDR compliance. Feeds CAPA and risk management update. |
| eQMS platform | (platform TBD) | Month 9 | Candidates: MasterControl, Greenlight Guru, SimplerQMS, Veeva Vault QMS. Must be 21 CFR Part 11 compliant with DHF and IEC 62304 modules. |
| VP Quality / Quality Lead hire | (org) | Month 6 | Interim authority: CEO. First action: gap assessment vs 21 CFR Part 820 + ISO 13485:2016 with external QMS consultant (Month 3). |
| Safety MCU unit-level FMEA | NP-SW-001 §11 | TBD | IEC 62304 §7.1 for Class C: identify all SW-01 failure modes (SW01-M01..M08) with potential harm and response time. |
| Software Bill of Materials (SBOM) | NP-SBOM-001 | Year 2 | Required for 510(k) cybersecurity submission. |

# Section 5 — T2 Modalities Firmware-Specified

## 5.1 sLORETA-Guided HD-tDCS (NP-FW-HD-001 Rev A, firmware/sloreta_hdtdcs/)

Personalised targeting for T2 depression, DLPFC modulation, and neuromodulation protocols. Software BASELINED (G3-07). Hardware FAI PENDING — requires T2 prototype.

| Parameter | Specification |
| --- | --- |
| sLORETA weight matrix | W: 2,447 voxels × 21 channels (≈205KB). Precomputed offline on 7mm MNI grid. Loaded into LPSDR4 at session start. |
| On-device runtime — covariance | 64-epoch accumulation → C (21×21 symmetric, 1.8KB SRAM) |
| On-device runtime — source power | P[v] = W_v^T × C × W_v for all 2,447 voxels → (9.6KB LPSDR4) → peak voxel detection |
| Electrode mapping | Jurcak et al. 2007 MNI scalp atlas for 21-ch T2 cap. E_center = nearest electrode to MNI target. 4 cathodes = nearest neighbours with ≥2 quadrant angular spread constraint. |
| Current distribution | Charge-balanced: anode I, each cathode −I/4. 30s ramp up/down, 100ms tick. |
| Safety — charge density | Software monitor at 95% of 40µC/cm² → abort. Safety MCU owns enable GPIO with 1kHz AC impedance check at session start. |
| Per-electrode limits | ≤2mA per electrode; focal electrode density ≤6 A/m² (Bikson lab safety limits for 3.5mm electrode geometry) |
| Concurrent EEG | 200µs blanking on current step edges. ≥20dB SNR specification. |
| Clinical targets (7) | DLPFC_L, DLPFC_R, VLPFC_L, ACC, MPFC, M1_L, M1_R. Bilateral 4×1 and standard 2-electrode fallback also implemented. |
| BOM delta | $0 — Ag/AgCl dual-rated electrodes already in T2 cap specification; tACS driver is existing T2 hardware |
| Evidence | Jog/UCLA 2025 (n=71, personalised MRI-guided HD-tDCS, significant depression improvement + grey matter changes); BRIGhTMIND 2024 (n=255, connectivity-guided iTBS, personalised targeting outperforms fixed F3) |

| FAI Item | Test | Status |
| --- | --- | --- |
| FAI-HD01 | Electrode localisation accuracy ≤15mm on anatomical phantom | PENDING — T2 prototype required |
| FAI-HD02 | Electrode mapping software correctness (all 6 criteria) | SOFTWARE PASS |
| FAI-HD03 | Spatial focality FWHM ≤25mm on phantom | PENDING — T2 prototype required |
| FAI-HD04 | Concurrent EEG SNR ≥20dB on bench | PENDING — T2 prototype required |

G3-07 SOFTWARE BASELINED. Full gate closure requires hardware FAI on T2 prototype.

## 5.2 Cervical VNS Safety Interlock (NP-FW-CVNS-001 Rev A, firmware/cervical_vns/)

Dual-processor cardiac interlock for carotid-proximity tcVNS stimulation. RISK-25 MITIGATED. G3-08 SOFTWARE BASELINED.

| Parameter | Specification |
| --- | --- |
| GPIO ownership | Safety MCU (STM32G071) owns CVNS_ENABLE_L and CVNS_ENABLE_R GPIO independently of main processor |
| R-peak signalling | Main processor Pan-Tompkins PPG R-peak detection drives RPEAK_IN GPIO (5ms active-high pulse) to safety MCU |
| HR monitoring | Safety MCU TIM2 (1MHz) measures R-R intervals → instantaneous HR → rolling 5s window vs pre-stimulation baseline |
| Cutoff trigger | HR change >15 BPM within 5s observation window |
| Cutoff latency | TIM6 ISR (5ms tick) → GPIO LOW within <5.1ms worst-case (specification: <100ms — 20× margin) |
| Baseline cross-validation | Main processor PPG-derived HR and safety MCU GPIO-timer-derived HR must agree within 5 BPM before CVNS_ENABLE granted |
| Re-enable requirements | All three conditions satisfied: 30s lockout elapsed + explicit app confirmation + repeat impedance check + baseline re-established |
| Stimulation waveform | Biphasic charge-balanced; 10s ramp-up, 5s ramp-down via np_cvns_stim |
| UHDR record | Full RR series, baseline HR, HR at cutoff event, timestamps, raw impedance readings |
| SHDR record | Cutoff flag only (no HR values), impedance pass/fail, fault count — no user biology |
| BOM delta | +$35–55 for complete cervical tcVNS accessory module (neck-worn, gel electrode assembly, consumable gel pad 5-pack) |

## 5.3 Cervical VNS 510(k) Q-Sub Strategy (NP-REG-CVNS-001 Rev A)

| Item | Detail |
| --- | --- |
| Predicate devices | electroCore gammaCore K163334 (cluster headache, 2016) + K173323 (migraine, 2017) |
| Substantial equivalence basis | Same intended use (cluster headache + migraine) + same/safer technology (biphasic charge-balanced, ≤2mA vs predicate ≤24mA; NeuroPulse 12× lower current) |
| Key differences — mitigated | Lower current: more conservative than predicate (SE positive). Bilateral option: ≤2mA supported by literature review + bench. Cardiac interlock: safety enhancement beyond predicate — SE positive. |
| Standards | IEC 60601-1, IEC 60601-2-10, IEC 62304 Class C (safety MCU) + Class B (main), IEC 62366-1, 21 CFR 820 |
| Q-Sub questions | 6 questions drafted for FDA CDRH DNPMD. Type B meeting target: Month 20 from T1 launch. |
| Clinical evidence | FAI-CV03 tolerability (3 healthy adults, IRB required) + predicate clinical programme + literature review |
| Extended indications | Depression, PTSD, post-stroke rehabilitation: planned for subsequent 510(k) via T2 seeding studies |

# Section 6 — Tooling Status Update

| Tooling Document | Number | Status | Key Open Items |
| --- | --- | --- | --- |
| Zone module tooling | NP-TOOL-ZM-001 Rev A | RELEASED | NP-COORD-001 mould design review checklist (all 8 features F-01 through F-08) must be signed before steel cut. LED emitter pulse current verification pending. |
| Smart zone module tooling variant | NP-TOOL-ZM-SM-001 Rev A | RELEASED | Mechanical key +2.0mm lateral — prevents mis-insertion into unmodified hub. 5-item DRC checklist. Dependent on base ZM-001 tooling release. |
| Shell tooling | NP-TOOL-SHELL-001 Rev A | RELEASED | CFRP Ra ≤1.6µm written supplier confirmation (RISK-20 OPEN — BLOCKING for tooling release). Hub tooling sub-items (probe dock, large-radius Boa channel, tool-free fan, anchor posts) pending. |
| Lens / goggle assembly tooling | NP-TOOL-LENS-001 Rev B | RELEASED | AgNW coating supplier qualification BLOCKING (P-01). N42 magnet ≥1mm polymer wall CAD verification required. OI-10 CAD sign-off for F-07 tether anchor (G2-13 OPEN). |
| Hub PCB | NP-HW-HUB-001 Rev B | SPECIFIED | Hub PCB Rev B Gerber build (OI-HUB-03) required — BLOCKING for FAI-SM-04 and FAI-SM-06 optical bench tests. |
| FPC layout — base module | NP-HW-FPC-001 Rev D | FROZEN | PD2 at X=33.0mm, Y=39.0mm locked. OI-3 CLOSED. No changes permitted without design change order. |
| FPC layout — smart module | NP-HW-FPC-001 Rev E (NP-FPC-ZM-SM-01) | FROZEN | 550 LED tri-wavelength layout locked. Rigidizer sub-board 22×14mm confirmed. Dependent on hub PCB Rev B Gerber for optical FAI. |

# Section 7 — Provisional Accessories

PROVISIONAL items are not committed. They are described here to document the provisioning decisions already made in tooling and firmware that enable them at zero incremental cost.

## 7.1 40Hz Vibrotactile Mastoid LRA Pad

Status: PROVISIONAL — release contingent on Cognito Therapeutics HOPE Phase 3 results (n=670, mid-2026). If positive → release within 6 months. Hardware anchor boss provisioned in NP-TOOL-SHELL-001 Rev A F-04 (temporal wing) at zero incremental tooling cost.

| Parameter | Specification |
| --- | --- |
| Actuator | LRA 8–10mm diameter (e.g. Jinlong JMC0834 or equivalent). LRA preferred over ERM: precise 40Hz frequency control, low harmonic distortion, flat resonance profile at 40Hz drive. |
| Driver IC | Texas Instruments DRV2605L (I2C, open-loop mode). 40Hz ± 0.5Hz locked. Amplitude control via gain register. No resonant frequency tuning required. |
| Placement | Posterior temporal / mastoid process. Direct bone coupling to skull. Adjacent to temporal lobe somatosensory representation. Compatible with existing temporal stability wing anchor boss (F-04 in NP-TOOL-SHELL-001 Rev A). |
| Output specification | 0.6–1.2G peak acceleration at skin surface. 40Hz ± 0.5Hz. 100% duty cycle for 20-minute session target. |
| Form factor | 30mm diameter silicone-overmoulded pad. Shore 30A silicone contact face. Clip attachment to temporal wing. 3-pin pogo or JST connector to hub accessory port. |
| Power | ~80–120mW continuous from hub accessory port (existing 500mA capability). No device battery drain. No Apple Watch battery impact. |
| Firmware | 40Hz square-wave drive pattern. Session start/stop synchronised with audio/visual 40Hz channels via hub. Amplitude ramp 2s up/down (comfort). Hub module ID: NP_MOD_VIBROTACTILE (0x10) — stub already registered in firmware/hub_control/. |
| BOM | $4–7 per pad (LRA $1.50–2.50 + DRV2605L $1.20–1.80 + PCB/passives $0.50 + silicone housing $0.80–1.20). Bilateral pair: $8–14. |
| Retail pricing | $49–79 per pad / $79–119 bilateral pair |
| vs Apple Watch Taptic Engine | Mastoid bone coupling vs wrist-to-brain soft tissue attenuation. 40Hz ± 0.5Hz precision vs uncharacterised Taptic Engine output. Hub-powered vs Watch battery. Use Watch app as supplement; mastoid pad for full protocol. |
| Decision gate | HOPE Phase 3 positive → release within 6 months. Negative → temporal wing boss retained in tooling (zero cost, zero obligation). |

## 7.2 Apple Watch Sync App (Provisional)

Companion watchOS app extending session experience to Apple Watch. BOM delta $0. Software development cost only. Regulatory position: all Watch functions declared session monitoring / UI aids — not therapeutic delivery. No therapeutic claims attach to Watch-delivered functions.

| Sync Channel | Function | Notes |
| --- | --- | --- |
| Channel 1 — Haptic sync | Core Haptics delivers 40Hz pattern synchronised with hub session clock. Wrist somatosensory channel on top of mastoid LRA pad. | Not a standalone therapeutic. App caveat: 'Works best with NeuroPulse mastoid vibrotactile accessory.' Pending OI-WA-01 thermal characterisation. |
| Channel 2 — Audio sync | Watch plays binaural beats / isochronic tones / breathing pacer via AirPods via WatchConnectivity, synchronised to hub session. | Useful when bone conduction is reserved for breathing pacer cue. Phase 2 development (Month 3–4). |
| Channel 3 — Visual sync | Watch display shows 40Hz visual flicker or EMDR L/R indicator arrow, synchronised to goggle session. Also: coherence score live feed, HRV breathing ring, session status. | Phase 4 blocked on OI-WA-02: screen brightness ≥100 nits at 40Hz characterisation. Phase 1 includes session status and HRV breathing ring only. |

# Section 8 — iOS App Development Roadmap (NP-APP-ROADMAP-001 Rev A)

Core iOS app scope (Year 1 launch target):

BLE GATT custom service for hub connection (hub antennas in control hub, not headset)

Session setup: protocol selection from 19-protocol predefined library or NPPS custom file import

Real-time session display: per-modality status, per-zone irradiance and dose bars, EEG band power, HRV coherence score

UHDR review: EEG waveforms, HRV time series, neurofeedback session history, 30-session trend graphs

Consumable management: impedance-triggered hydrogel tip prompts, sleeve count tracking, one-tap reorder links

Firmware OTA management: Bank B staging, Ed25519 signature verification display, user confirmation workflow

Clinical consent engine: four-screen a priori research consent, per-study notification delivery, research portal access

SHDR maintenance alerts: LED aging ratios, fouling flags, predictive maintenance with one-tap consumable ordering

Power display: 'Power level: reduced' when non-PD-compliant charger detected — informative, never blocking (EU compliance)

Apple Watch sync app — Year 1 post-launch delivery (4 phases):

| Phase | Features | Target Month | Blockers |
| --- | --- | --- | --- |
| Phase 1 | Session status display (protocol name, elapsed time, modalities active). HRV breathing ring (expanding/contracting ring synchronised to pacer). Haptic end-of-session alert. | Month 1–2 development App Store target Month 3 | None |
| Phase 2 | Audio sync: binaural beats / isochronic tones / breathing pacer delivered to AirPods via WatchConnectivity framework, synchronised to hub session clock. | Month 3–4 | OI-WA-01 thermal characterisation of Watch during continuous audio transmission |
| Phase 3 | Haptic sync: Core Haptics 40Hz pattern synchronised to hub session clock. Sub-50ms latency target via BLE GATT custom service. | Month 4–5 | OI-WA-01 Watch thermal characterisation (continuous haptics 20 min) |
| Phase 4 | 40Hz visual flicker on Watch display (GENUS-compatible). EMDR L/R indicator arrow synchronised to goggle session. | Month 6+ | OI-WA-02: Watch screen brightness ≥100 nits at 40Hz flicker characterisation required before clinical protocol claim |

# Section 9 — Risk Register Update

25 risks total (RISK-01 through RISK-25). 23 MITIGATED. 2 OPEN (external actions required). All risks formally entered under ISO 14971:2019 change control via NP-RM-001 Rev A.

## 9.1 New Risk Added Since Revision 4

| Risk | Hazard | Control | Status |
| --- | --- | --- | --- |
| RISK-25 | Cervical VNS: cardiac reflex (vasovagal syncope / bradycardia / vagally-mediated arrhythmia) from stimulation near carotid sheath / vagus trunk | Dual-processor cardiac interlock: Safety MCU owns CVNS_ENABLE GPIO. R-peak GPIO from main processor. TIM6 ISR asserts GPIO LOW within <5.1ms of HR change >15 BPM in 5s window. Baseline cross-validation required before enable. 30s re-enable lockout. Firmware: NP-FW-CVNS-001 Rev A. | MITIGATED — firmware interlock baselined (G3-08 SOFTWARE BASELINED). FAI-CV02 hardware bench timing verification PENDING (oscilloscope + R-peak signal generator — T2 prototype required). FAI-CV03 (3 healthy adults, IRB required). |

## 9.2 Risks Still OPEN

| Risk | Issue | Status | Action required |
| --- | --- | --- | --- |
| RISK-03 | 400 mW/cm² peak pulsed irradiance claim — no regulatory opinion obtained. Scope now expanded to include: 1064nm LED array FDA pathway; aggregate irradiance ceiling 600 mW/cm² (OI-PBM-05); T2 combined 1064+1170nm session dual-standard hazard; depth-tier penetration FTC substantiation. | OPEN — BLOCKING for all public irradiance and depth-tier claims | NP-REG-PBM1064-001 Rev A submitted to outside regulatory counsel (12 questions Q1–Q12). Opinion letter required. Estimated $8,000–25,000 total with scope expansion. No timeline from counsel yet. |
| RISK-20 | CFRP shell slot rim surface finish Ra ≤1.6µm — required for self-sealing gasket seating. Unconfirmed from CFRP shell tooling supplier. | OPEN — BLOCKING for tooling release | Obtain written confirmation from CFRP shell tooling supplier: letter + Ra measurement data from representative coupon. NP-PROC-SUP-001 supplier qualification items SUP-M-07 and SUP-B-01. If Ra > 1.6µm unavoidable: escalate to ME + EE; gasket or shell geometry revision required. |

RISK-04 (PDMS CAT-C supplier) remains a separate blocking item — not a risk register entry but a supplier selection gate (NP-PROC-SUP-001 SUP-C-08). 200-cycle IEC 60068-2-14 thermal cycling qualification required before any production FPCs are built.

# Section 10 — Open Items and Gate Status

## 10.1 Engineering Coordination Checklist Gate Items (NP-COORD-001)

| Gate Item | Description | Status | Document |
| --- | --- | --- | --- |
| G1-15 | FPC layout freeze with PD2 position locked | CLOSED | NP-HW-FPC-001 Rev D |
| G1-16 | eMMC partition architecture + dual-bank OTA bootloader spec | CLOSED | NP-FW-EMMC-001 Rev A |
| G2-10 | Zone bone conduction announcement firmware | CLOSED | NP-FW-ZA-001 Rev A (firmware/zone_announce/) |
| G2-11 | EEG cable routing path in shell CAD model | CLOSED | NP-DRV-SHELL-001 Rev B |
| G2-12 | HRV biofeedback firmware | CLOSED | NP-FW-HRV-001 Rev A (firmware/hrv_biofeedback/) |
| G2-13 | Lens assembly G2 deliverable | OPEN | NP-TOOL-LENS-001 Rev B — pending OI-10 CAD sign-off for F-07 tether anchor |
| G3-07 | sLORETA-guided HD-tDCS | SOFTWARE BASELINED | NP-FW-HD-001 Rev A — FAI-HD01/03/04 hardware PENDING (T2 prototype) |
| G3-08 | Cervical VNS safety interlock | SOFTWARE BASELINED | NP-FW-CVNS-001 Rev A — FAI-CV01/02/03 hardware PENDING (T2 prototype + IRB) |

## 10.2 Pending Decisions — Must Resolve Before Tooling Cut

BLOCKING for tooling (immediate action required):

Product name trademark clearance — US, EU, Canada, Australia. Required before any external conversation.

400 mW/cm² regulatory opinion letter (RISK-03 BLOCKING for all public irradiance claims). Counsel engaged via NP-REG-PBM1064-001 Rev A.

CFRP shell slot rim Ra ≤1.6µm — written supplier confirmation (RISK-20 BLOCKING for tooling release). NP-PROC-SUP-001 SUP-M-07/SUP-B-01.

PDMS CAT-C supplier selection + 200-cycle IEC 60068-2-14 thermal cycling qualification (RISK-04 BLOCKING for production start). Supplier lead time 8 weeks from selection.

Zone module mould design review — NP-TOOL-ZM-001 §5 checklist (all 8 features F-01 through F-08) signed off before steel cut.

LED emitter pulse current rating verification (660nm + 808–830nm FPC candidates at 120–180mA per LED).

Hub PCB Rev B Gerber build (OI-HUB-03) — BLOCKING for FAI-SM-04, FAI-SM-06 optical bench tests.

Required before first firmware release:

Factory calibration procedure for InGaAs PD coefficients (K_PD1_1064, K_PD2_1064, K_ratio_nom): integrating sphere reference at 1064nm; write to Config partition at manufacture (OI-PBM-04 in NP-FW-PBM1064-001).

ZONE_ID firmware debounce spec written into firmware requirements document (3×ADC at 100ms, ≥2/3 pass) — confirm in NP-FW-PBM1064-001.

QMS milestones:

Hire Quality Lead (VP Quality) — Month 6 target; interim authority CEO. First action: gap assessment vs 21 CFR Part 820 + ISO 13485:2016 (Month 3).

Design and Development Plan NP-DP-001 — Month 3.

Design Input/Output Traceability Matrix NP-DT-001 — Month 6.

Human Factors Engineering Plan NP-HFE-001 (IEC 62366-1) — Month 9.

Post-Market Surveillance Plan NP-PMS-001 — Month 12 (before T1 launch).

eQMS platform selection — Month 9.

Safety MCU unit-level FMEA (NP-SW-001 §11, IEC 62304 §7.1 Class C) — SW01-M01..M08.

Clinical and commercial milestones:

Hub tooling: probe dock + anchor posts + large-radius Boa cable channel + tool-free fan (quarter-turn captive fastener).

HFE formative study for sliding eject lever — 5 subjects with Parkinson's/post-stroke grip weakness (OI-4 in NP-TOOL-ZM-001, FAI-A15 in NP-FAI-ZM-001).

SAB formation — Tsai outreach first (617-324-0305, SAB role only; Cognito Therapeutics conflict prevents PI role). Budget $50,000–80,000/year for 5-person SAB.

First researcher contacts: Rashidi-Ranjbar (neda.rashidi-ranjbar@unityhealth.to) → Jog (mjog@mednet.ucla.edu) → Naeser (mnaeser@bu.edu).

SBIR Phase I application — commission at company formation.

Partner optician network contract (S3 Rx clip programme — Tier A service network).

FAI-CV01 electrode placement bench (anatomical phantom — T2 prototype required).

FAI-CV02 cardiac interlock timing bench (oscilloscope + R-peak signal generator — T2 prototype required).

FAI-CV03 cervical VNS tolerability (3 healthy adults — IRB approval required).

# Section 11 — Completed and Locked Decisions (Comprehensive)

All items below are LOCKED. Changes require a formal design change order under NP-QMS-DC-001. Items are grouped by category. Items marked (NEW Rev 5) were locked since Revision 4.

## 11.1 Product Strategy and Commercial

Two-tier product strategy with shared chassis, processor stack, app, and USB-C connectivity (T1 Home / T2 Pro)

No mandatory subscription — all core functions offline-capable permanently

UHDR/SHDR terminology framework — complete separation, user data never accessed by NeuroPulse

Clinical consent engine: use case library → minimum necessary data mapper → plain-language consent document

Research partnership: 4-layer a priori consent, results opt-in, suggestion portal, crowdfunding catalyst

POA workflow with human document review (3 business days, annual re-verification, jurisdiction flagging)

Service network: Tier A optical + Tier B electronics + Tier C retail + depot backstop

Charger policy: auto-include correct charger at every upgrade + $19 upfront 65W option at checkout as intent signal

Consumables: intranasal sleeves (only authenticated consumable), hydrogel tips, VNS pads — measurement-triggered prompts

Predictive maintenance via SHDR fleet telemetry — three phases (population-average → fleet LSTM → Bayesian personalisation)

On-device research data anonymisation (k≥10): fresh per study, NeuroPulse never holds raw UHDR

Consent withdrawal immediately effective for all future data flows from any time period (including historical sessions)

## 11.2 Hardware Platform

Main processor: NXP i.MX RT1062 (Cortex-M7, 600MHz, FPU+DSP+SIMD, 1MB on-chip SRAM + 32MB LPSDR4)

Safety MCU: STM32G071 (NOT G031 — G031 has only 8KB SRAM, insufficient for EMF firmware). 64MHz, 36KB SRAM, 128KB flash, bare-metal

Storage: 8GB industrial eMMC (SLC cache, 30,000+ P/E cycles), LittleFS filesystem

Connectivity: USB-C 3.2 Gen1 (default, zero RF, <1ms) + BT 5.3 LE Audio + Wi-Fi 6; antennas in control hub NOT headset; single rear toggle

Safety MCU physically owns all stimulation enable GPIO — app crash cannot cause unsafe stimulation

SPI heartbeat every 200ms; 1,500ms watchdog → all-stimulation cutoff <50ms

22F supercapacitor (absorbs LED duty-cycle transients; allows 50% aging over 5 years)

5-layer passive EMF shielding: CFRP outer + mu-metal liner + palladium-coated polyester + EMI absorber foam + USB-C port filters

Active Helmholtz shielding: 3-axis fluxgate magnetometers + Helmholtz coil pairs (35–45dB ELF magnetic combined)

No NFC anywhere — optical code + resistive pogo pin authentication for consumables

N42 (not N52) neodymium magnets in lens rim — better impact tolerance, −$0.80 BOM

AgNW (silver nanowire) outer conductive lens coating — replaces ITO (ITO 0.5% strain-to-failure; AgNW ≥5%)

Hard clamshell case (not soft pouch) — lens scratch protection, probe dock, doubles as shipping container

Braided aramid USB-C cable with dual silicone strain relief — 50,000+ flex cycle rating; spare in box

Palladium-coated EMF shielding fabric — tarnish-immune for device lifetime; fleet SHDR monitors attenuation

ADS1299 internal reference self-calibration at every session start (eliminates EEG gain/offset drift)

Spring-decoupled electrode pods: 80–120g contact force, ±12mm travel, independent of dial tension

Boa-style occipital dial, 10cm range, 0.5mm/click, 50,000-cycle rated, enclosed PTFE-lined cable channel

1 adult SKU covers 52–62cm heads

## 11.3 FPC Zone Module Engineering Package

600 total LEDs per headset from day one: 300×660nm + 300×808–830nm on FPC strips (no Rev A/Rev B LED count ramp)

FPC 20-pin pinout locked: Pin 18 = ZONE_ID (resistor per zone: ZM-01=10kΩ, ZM-02=22kΩ, ZM-03=47kΩ, ZM-04=100kΩ, ZM-05=220kΩ, all 1%); Pin 19 = PD2_CATHODE

FPC layout frozen: PD2 at X=33.0mm, Y=39.0mm from module reference corner. 1.6mm annular ring, scalp-facing hard gold ≥0.5µm. OI-3 CLOSED (NP-HW-FPC-001 Rev D)

Dual photodiode dose-metering (RISK-14 Option B): PD2 on scalp-facing PDMS surface (pin 19); PD1/PD2 ratio separates fouling from LED aging; T1/T2 identical zone module mould

Five-layer zone module keying (RISK-15): mechanical key + ZONE_ID resistor + ISO 17049 braille/numeral + shell tactile dots + bone conduction audio; covers colour-blind and blind users

Self-sealing co-moulded silicone gasket (RISK-16 Option A): Shore 40–50A medical silicone, D-section 2.5×2.0mm, 20% compression when seated. No user RTV required. IPX4 rated after 10 swap cycles (FAI-IPX-02 BLOCKING)

Sliding eject lever (RISK-22 Option A): 10–12mm lever arm, 3:1 mechanical advantage, ≤1N extraction force. Accessibility target: Parkinson's Hoehn & Yahr II–III

ZONE_ID firmware debounce: 3× ADC reads at 100ms intervals; ≥2/3 must pass (RISK-18)

PDMS bonding process: 75nm RF magnetron sputtered SiO₂ on PI surface + O₂ plasma activation. ≥150 N/m peel strength. 200-cycle IEC 60068-2-14 thermal cycling qualification BLOCKING before production (RISK-04)

Zone module connectors: Hirose FH34S-20S-0.5SH (or JAE FF03) — 0.5mm pitch, back-flip lever ZIF, ≥1,000 insertion cycles. 0.35mm pitch insufficient. Must be specified before PCB layout.

Plasma-activated anti-fouling PDMS optical windows — PDMS–PI bond via 75nm SiO₂ interlayer + O₂ plasma; 174–860 N/m peel force

EEG cable routing (RISK-21, OI-09 CLOSED): 8×5mm main trunk + 6×4mm lateral branches on outer CFRP inner surface. Hub → occipital arch (~120mm) → crown junction → left/right branches. 18–22mm FPC-to-EEG separation. DRC-18/22/23 CLOSED

Multi-FPC bundle management (RISK-17): ≥2mm inter-FPC separation; ≥15mm FPC-to-EEG; 3 anchor bosses per FPC; all 5 hub ZIF connectors on same PCB edge

Interface protection covers (all tethered): zone slot plugs ×5 (Shore 30A silicone, 5 colours), accessory port covers ×3 (magnetic steel disc), lens rim guards ×2 (Shore 85A TPU). Anchor posts moulded into shell at zero cost if specified before first cut.

## 11.4 New Decisions Locked in Revision 5

(NEW Rev 5) 1064nm smart zone module — Path B architecture locked: ATtiny402 I2C slave + 3× IRLML6344 N-FETs. 550 LEDs (200×660nm + 200×808nm + 150×1064nm). InGaAs PD Hamamatsu G12180-010A. EPITEX L1064-02AU via Marubeni. BOM delta +$23–28/module. Retail $149–199/zone. (NP-HW-FPC-001 Rev E)

(NEW Rev 5) Smart module ZONE_ID = 3.3kΩ on pin 18 (ADC < 1100 counts, 12-bit) — distinct from base module 10kΩ–220kΩ range (ADC 1100–4000 counts)

(NEW Rev 5) Smart module mechanical key: +2.0mm lateral face width vs base module — prevents physical insertion into hub slots without Rev B TIA gain switch. Safety mechanical lock.

(NEW Rev 5) Hub PCB Rev B architecture SPECIFIED: Vishay DG2788A per-slot TIA gain switch + NXP PCA9546A 4-ch I2C isolation switch. Hub BOM delta $1.41–1.91. Sequencing: ZONE_ID debounce → GAIN_SEL assert → I2C mux enable. (NP-HW-HUB-001 Rev B)

(NEW Rev 5) eMMC partition architecture BASELINED: 9-partition GPT, AES-256-XTS UHDR (Argon2id biometric key — NeuroPulse never holds), AES-256-XTS SHDR (manufacturing key), LittleFS on UHDR/SHDR/Config. 27-element session data classification table. (NP-FW-EMMC-001 Rev A, G1-16 CLOSED)

(NEW Rev 5) Dual-bank OTA bootloader WRITTEN: Ed25519 self-contained verify, 9-step OTA, DFU 1.1 recovery, SNVS_LPGPR0 bank flag, streaming SHA-256. (firmware/bootloader/)

(NEW Rev 5) Hub control program WRITTEN: 4 FreeRTOS tasks (module probe, protocol runner, telemetry, safety SPI heartbeat). Ed25519 protocol verification. np_module_registry. T2 stubs registered. (firmware/hub_control/, NP-FW-HUB-001 Rev A)

(NEW Rev 5) NPPS protocol scripting language RELEASED: .npps format, 15 modality blocks, composite/interval/limits blocks, 19 predefined protocols, app-side compiler hubCompiler.ts. (docs/npps-reference.md)

(NEW Rev 5) Hub wire format expanded: 6 T2 module type IDs registered. mods_active_mask uint8_t → uint32_t.

(NEW Rev 5) Zone bone conduction announcement firmware WRITTEN: RISK-15 Layer 5, 5 slot state machines, DDS audio, 256-entry Q15 sine LUT, SAI3 I2S, SHDR auth log. (firmware/zone_announce/, G2-10 CLOSED)

(NEW Rev 5) HRV biofeedback firmware WRITTEN: 6-module library (ppg, coherence 256-pt Welch PSD, pacer, tavns_sync, eeg_biofeedback, session). Four protocols: standalone, HRV+taVNS synchronised, HRV+EEG dual biofeedback, HRV+PBM. (firmware/hrv_biofeedback/, G2-12 CLOSED)

(NEW Rev 5) sLORETA HD-tDCS firmware SOFTWARE BASELINED: scalar sLORETA W matrix (2447×21, ≈205KB LPSDR4), 4×1 ring montage, 7 clinical targets, 30s ramp, 200µs EEG blanking. (firmware/sloreta_hdtdcs/, G3-07 SOFTWARE BASELINED)

(NEW Rev 5) Cervical VNS firmware SOFTWARE BASELINED: dual-processor cardiac interlock, TIM6 ISR <5.1ms GPIO cutoff, baseline cross-validation, 3 modules. RISK-25 MITIGATED. (firmware/cervical_vns/, G3-08 SOFTWARE BASELINED)

(NEW Rev 5) Cervical VNS 510(k) Q-Sub BASELINED: electroCore gammaCore K163334+K173323 predicate, SE argument, Type B meeting Month 20. (NP-REG-CVNS-001 Rev A)

(NEW Rev 5) QMS ESTABLISHED: 6 documents released 2026-05-13. IEC 62304 classes locked (Safety MCU = Class C, Main processor = Class B, App = Class B). RISK-01..25 under ISO 14971 change control. DHF baseline established. (NP-QMS-001, NP-DHF-001, NP-QMS-DC-001, NP-RM-001, NP-SW-001, NP-QMS-CAPA-001)

(NEW Rev 5) Shell tooling specification RELEASED: F-01 zone plug anchor posts (×5 colour-coded), F-02 port cover anchor posts (×3), F-03 EEG cable routing channel (8×5mm), F-04 temporal wing boss (PROVISIONAL). (NP-TOOL-SHELL-001 Rev A)

(NEW Rev 5) Lens tooling Rev B RELEASED: F-07 goggle arm anchor hook (4mm ID loop, superior-posterior corner), shade clearance ≥4mm MR-19, tether ≤45mm MR-20 no FOV crossing. OI-10 tracks CAD sign-off. (NP-TOOL-LENS-001 Rev B)

(NEW Rev 5) FPC layout frozen: PD2 at X=33.0mm, Y=39.0mm — OI-3 CLOSED, G1-15 CLOSED. (NP-HW-FPC-001 Rev D)

(NEW Rev 5) T2 combined 1064+1170nm session protocol LOCKED: three-tier penetration, thermal throttle priority (1170nm first → CH_C → CH_B). (NP-SES-1064-001 Rev A)

(NEW Rev 5) RISK-03 scope expansion SUBMITTED: NP-REG-PBM1064-001 Rev A brief to counsel, 12 questions Q1–Q12, marketing gate on all 1064nm/aggregate irradiance/depth-tier claims until opinion letter received.

(NEW Rev 5) 1064nm bibliography addendum ADDED: NP-BIB-1064-001 Rev A, 6 entries including Yao et al. Science Advances 2022 (n=20 RCT, right PFC 1064nm, CDA working memory improvement, primary evidence).

(NEW Rev 5) Risk register: 25 risks (RISK-01..25), 23 MITIGATED, RISK-03 and RISK-20 OPEN, RISK-25 MITIGATED pending hardware bench.

(NEW Rev 5) iOS app roadmap RELEASED: NP-APP-ROADMAP-001 Rev A. Apple Watch sync app 4-phase Year 1 post-launch. Core app scope defined. BOM delta $0 for Watch app. (docs/neuropulse_ios_app_roadmap_001.md)

(NEW Rev 5) 40Hz vibrotactile mastoid LRA pad: PROVISIONAL pending HOPE Phase 3. Temporal wing anchor boss in NP-TOOL-SHELL-001 Rev A F-04 at zero incremental cost. Hub module stub registered.

(NEW Rev 5) Apple Watch sync app: PROVISIONAL. 3 sync channels declared monitoring/UI aids. BOM delta $0.

(NEW Rev 5) EEG cable routing CLOSED: OI-09 CLOSED, 8×5mm trunk, 18–22mm FPC separation, DRC-18/22/23 CLOSED, G2-11 CLOSED. (NP-DRV-SHELL-001 Rev B)

## 11.5 T2 Modality Decisions (all locked)

21-channel qEEG wet gel: full 10-20 + FC3/FC4 (M1 TMS targeting) + Oz (photoparoxysmal) + A1/A2 (linked-ear normative reference, on VNS clips)

TMS focal figure-8 coil: 0.1–0.5T, rTMS + TBS, non-conductive CFRP window, EMF cancellation gated off 5ms pre-pulse + 50ms post-pulse hold

1170nm deep PBM: laser diodes, 35–40mm subcortical depth, TEC stabilisation, ≤1,000 mW/cm²

Clinical tACS: ≤4mA, 16-channel arbitrary waveform

sLORETA-guided HD-tDCS: scalar sLORETA, 4×1 ring, 7 clinical targets, BOM delta $0, G3-07 SOFTWARE BASELINED

Cervical tcVNS (T2 accessory): carotid sheath, dual-processor cardiac interlock, 510(k) predicate electroCore K163334+K173323, BOM delta +$35–55

HIPAA cloud + EHR: FHIR R4, multi-patient dashboard, LSL streaming, scripting API, anonymised session tag (clinic-held patient-to-tag mapping)

# Section 12 — Document Register

All NeuroPulse design programme documents. Firmware source directories are declared as design output records in NP-DHF-001 Rev A. Documents marked (NEW Rev 5) were created since Revision 4.

| Document Title | Number | Rev | Date | Location |
| --- | --- | --- | --- | --- |
| Design Brief Revision 1 | NP-DB-001 Rev 1 | R1 | 2026-Q1 | docs/neuropulse_design_brief.docx |
| Design Brief Revision 2 | NP-DB-001 Rev 2 | R2 | 2026-Q1 | docs/neuropulse_design_brief_r2.docx |
| Design Brief Revision 3 | NP-DB-001 Rev 3 | R3 | 2026-Q2 | docs/neuropulse_brief_r3.docx |
| Design Brief Revision 4 | NP-DB-001 Rev 4 | R4 | 2026-Q2 | docs/neuropulse_brief_r4.docx |
| Design Brief Revision 5 (this document) | NP-DB-001 Rev 5 | R5 | 2026-05-16 | docs/neuropulse_brief_r5.docx |
| Clinical Evidence Bibliography | NP-BIB-001 | A | 2026-Q2 | docs/neuropulse_bibliography.docx |
| Researcher Candidate List | NP-CLIN-RES-001 | A | 2026-Q2 | docs/neuropulse_researchers.docx |
| SBIR Phase I Draft | NP-SBIR-001 | A | 2026-Q2 | docs/neuropulse_sbir_phase1_draft.docx |
| FPC Zone Module Specification | NP-HW-FPC-001 | D/E | 2026-05-09/13 | docs/neuropulse_fpc_zone_module_spec_revA.docx |
| FPC Procurement Requirements | NP-PROC-FPC-001 | A | 2026-Q2 | docs/neuropulse_fpc_procurement_requirements.docx |
| Zone Module Risk Register | NP-RM-ZM-001 | A | 2026-Q2 | docs/neuropulse_fpc_zone_module_risks_revA.docx |
| Shell FPC Routing Review | NP-DRV-SHELL-001 | B | 2026-05-10 | docs/neuropulse_shell_fpc_routing_review.docx |
| FAI Zone Module Checklist | NP-FAI-ZM-001 | A | 2026-Q2 | docs/neuropulse_fai_zone_module.docx |
| Engineering Coordination Checklist | NP-COORD-001 | A.7 | 2026-05-11 | docs/neuropulse_eng_coordination_checklist.docx |
| Zone Module Tooling Specification | NP-TOOL-ZM-001 | A | 2026-Q2 | docs/neuropulse_tool_zone_module_001.docx |
| Supplier Selection Checklist | NP-PROC-SUP-001 | A | 2026-Q2 | docs/neuropulse_supplier_selection_checklist.docx |
| Clinical Trials Strategy | NP-CLIN-001 | A | 2026-Q2 | docs/neuropulse_clinical_trials_strategy.docx |
| Additional Modalities | NP-MOD-EXT-001 | A | 2026-Q2 | docs/neuropulse_additional_modalities.docx |
| eMMC Partition Architecture and Encryption (NEW Rev 5) | NP-FW-EMMC-001 | A | 2026-05-11 | docs/neuropulse_fw_emmc_001.docx |
| HRV Biofeedback Protocol Firmware Spec (NEW Rev 5) | NP-FW-HRV-001 | A | 2026-05-11 | docs/neuropulse_fw_hrv_001.md |
| Zone Bone Conduction Announcement Firmware Spec (NEW Rev 5) | NP-FW-ZA-001 | A | 2026-05-11 | docs/neuropulse_fw_zone_announce_001.md |
| sLORETA HD-tDCS Firmware Spec (NEW Rev 5) | NP-FW-HD-001 | A | 2026-05-11 | docs/neuropulse_fw_sloreta_hdtdcs_001.md |
| Cervical VNS Safety Interlock Firmware Spec (NEW Rev 5) | NP-FW-CVNS-001 | A | 2026-05-11 | docs/neuropulse_fw_cvns_001.md |
| Cervical VNS 510(k) Pre-Submission (NEW Rev 5) | NP-REG-CVNS-001 | A | 2026-05-11 | docs/neuropulse_cvns_510k_presub_001.md |
| QMS Manual (NEW Rev 5) | NP-QMS-001 | A | 2026-05-13 | docs/neuropulse_qms_manual_001.md |
| Design History File Index (NEW Rev 5) | NP-DHF-001 | A | 2026-05-13 | docs/neuropulse_dhf_index_001.md |
| Design Controls Procedure (NEW Rev 5) | NP-QMS-DC-001 | A | 2026-05-13 | docs/neuropulse_design_controls_001.md |
| ISO 14971 Risk Management Plan (NEW Rev 5) | NP-RM-001 | A | 2026-05-13 | docs/neuropulse_risk_mgmt_plan_001.md |
| IEC 62304 Software Development Plan (NEW Rev 5) | NP-SW-001 | A | 2026-05-13 | docs/neuropulse_sw_dev_plan_001.md |
| CAPA Procedure (NEW Rev 5) | NP-QMS-CAPA-001 | A | 2026-05-13 | docs/neuropulse_capa_001.md |
| Shell Tooling Specification (NEW Rev 5) | NP-TOOL-SHELL-001 | A | 2026-05-10 | docs/neuropulse_tool_shell_001.docx |
| Lens/Goggle Assembly Tooling Rev B (NEW Rev 5) | NP-TOOL-LENS-001 | B | 2026-05-10 | docs/neuropulse_tool_lens_001.docx |
| 1064nm Smart Module FPC Spec (NEW Rev 5) | NP-HW-FPC-001 Rev E / NP-FPC-ZM-SM-01 | E | 2026-05-13 | docs/neuropulse_hw_fpc_smart_001.md |
| 1064nm Smart Module FPC Procurement (NEW Rev 5) | NP-PROC-FPC-1064-001 | A | 2026-05-13 | docs/neuropulse_proc_fpc_1064_001.md |
| Smart Zone Module Tooling Variant (NEW Rev 5) | NP-TOOL-ZM-SM-001 | A | 2026-05-13 | docs/neuropulse_tool_zone_module_smart_001.md |
| Hub PCB Rev B (NEW Rev 5) | NP-HW-HUB-001 | B | 2026-05-13 | docs/neuropulse_hw_hub_pcb_revb_001.md |
| T2 Combined Session Protocol (NEW Rev 5) | NP-SES-1064-001 | A | 2026-05-12 | docs/neuropulse_session_protocol_1064_001.md |
| 1064nm PBM Bibliography Addendum (NEW Rev 5) | NP-BIB-1064-001 | A | 2026-05-13 | docs/neuropulse_bibliography_1064nm_001.md |
| RISK-03 Regulatory Opinion Scope Expansion (NEW Rev 5) | NP-REG-PBM1064-001 | A | 2026-05-13 | docs/neuropulse_reg_pbm1064_risk03_001.md |
| 1064nm Smart Zone Module Firmware Spec (NEW Rev 5) | NP-FW-PBM1064-001 | A | 2026-05-12 | docs/neuropulse_fw_pbm1064_001.md |
| iOS App Development Roadmap (NEW Rev 5) | NP-APP-ROADMAP-001 | A | 2026-05-11 | docs/neuropulse_ios_app_roadmap_001.md |
| NPPS Protocol Scripting Language Reference (NEW Rev 5) | NP-NPPS-001 | A | 2026-05-11 | docs/npps-reference.md |
| OTA Bootloader Firmware (NEW Rev 5) | NP-FW-EMMC-001 §8 (source) | A | 2026-05-11 | firmware/bootloader/ |
| HRV Biofeedback Firmware (NEW Rev 5) | NP-FW-HRV-001 (source) | A | 2026-05-11 | firmware/hrv_biofeedback/ |
| Zone Announcement Firmware (NEW Rev 5) | NP-FW-ZA-001 (source) | A | 2026-05-11 | firmware/zone_announce/ |
| sLORETA HD-tDCS Firmware (NEW Rev 5) | NP-FW-HD-001 (source) | A | 2026-05-11 | firmware/sloreta_hdtdcs/ |
| Cervical VNS Firmware (NEW Rev 5) | NP-FW-CVNS-001 (source) | A | 2026-05-11 | firmware/cervical_vns/ |
| 1064nm Smart Module Firmware (NEW Rev 5) | NP-FW-PBM1064-001 (source) | A | 2026-05-12 | firmware/pbm_1064nm/ |
| Hub Control Program (NEW Rev 5) | NP-FW-HUB-001 (source) | A | 2026-05-11 | firmware/hub_control/ |

NP-DB-001 Rev 5 — NeuroPulse Master Design Brief — May 16, 2026
CONFIDENTIAL — Pre-production design document. All decisions marked LOCKED are under QMS change control (NP-QMS-DC-001).
Decisions marked PROVISIONAL are contingent on external events as noted. Decisions marked PENDING require action before tooling cut.